## Instructions
Total Error vs Total Allowable Error
Laboratory | YOUR LABORATORY NAME/DEPARTMENT
Period under review: |  | From dd/mm/yy to dd/mm/yy |  |  |  |  | Signature |  | Lab Manager's signature
Instrument | Analyte | Control Sample | Lot number | Units | Target Value | # | TEa% | TEa (in units) | * | Mean | SD | TE (in units) | N | Pass/ Fail | Comment
Instrument name | Test name or reagent name | Control name and level | Current lot number | Units of measure | Target value for control as per package insert |  |  |  |  | Your mean for the period under review | Your SD for the period under review |  | The number of control measurements under reveiw |  | Your corrective action for any "FAIL"
Complete columns A to F with your lab information. This can be used as a template until you change lot numbers.
Gather your IQC data for a calendar month. Instruments should be able to give you your cumulative data for a set period. In Haematology (FBC), it might be easier to go per lot number.
In Column K, enter your cumulative mean.
In Column L, enter your cumulative SD.
In Column N, enter the number of control measurements used in the calculation on the mean & SD. If your "N" is less than 20, go back another month until you have a minimum of 20 results.
Column G is a code to reference where you got your target value from. Currently we are all going to use our Package Insert target values, so the number in there is "2". (See below the spreadsheet)
The YELLOW columns are locked. You may not change anything in these columns.
Once you have saved the information, Column O will automatically tell you if you passed or failed.
Investigate any levels of QC that FAIL, and record your corrective action.
Print out and get the lab manager to review & sign the print out. File with the Month's QC data.
The following sheets are for Chem, Haem & Coag.
Contact your QA Coordinator for assistance in setting it up specifically to your laboratory.

## Ref values (hidden)
|  | 1 | 2 | 3 | 4 |  |  | 6
|  | Biological Variation |  |  | PT Scheme Goals
|  | Optimal | Desirable | Minimum | RCPA |  | Thistle | CLIA | RIQAS
Matrix | Analyte | TEa | TEa | TEa | Lower Goal | Upper Goal | Allowable CV | TEa | (June2022)
S- | 11-Desoxycortisol |  | 27.1%
S- | 17-Hydroxyprogesterone |  | 29.7% |  | 2.0 nmol/L ≤ 10.0 nmol/L | 20% ≥ 10 nmol/L | 20%
U- | 4-hydroxy-3-methoximandelate (VMA) |  | 31.3%
S- | 5' Nucleotidase |  | 26.8%
U- | 5'-Hydroxyindolacetate, concentration |  | 26.5% |  |  |  | 20%
S- | α1-Acid Glycoprotein | 8.1% | 16.2% |  |  |  | 12%
S- | α1-Antichymotrypsin |  | 16.8%
S- | α1-Antitrypsin |  | 9.2% | 13.8% |  |  | 16% | 3sd
S- | α1-Globulins |  | 15.7%
U- | α1-Microglobulin, concentration, first morning |  | 43.9%
S- | α2-Globulins |  | 12.6%
S- | α2-Macroglobulin |  | 7.56% | 11.3% |  |  | 12%
U- | α2-Microglobulin output, first morning |  | 34.7%
P- | α-aminobutryic acid |  | 30.5%
S- | α-Amylase | 7.3% | 14.6% | 21.9% | 10 U/L ≤ 100 U/L | 10% ≥ 100 U/L | 19% | 30%
S- | α-Amylase (pancreatic) | 51.9% | 17.7%
U- | α-Amylase (pancreatic) | 44.4% | 88.82%
U- | α-Amylase concentration, random |  | 103.7% |  |  |  | 19%
P- | α-Carotene |  | 37.1%
S- | α-Carotene |  | 59.8%
S- | α-Fetoprotein (AFP) | 10.9% | 21.9% |  | 2 kIU/L ≤ 17 kIU/L | 12% ≥ 17kIU/L | 20% | 3sd | 13.3%
S- | α-Tocopherol |  | 16.5% | 24.7%
S- | Acetaminophen (Paracetamol) |  |  |  | 20 umol/L ≤ 200 umol/L | 10% ≥ 200 umol/L | 12%
S- | Acid phosphatase |  | 10.3% |  |  |  | 12%
S- | Acid phosphatase tartrate-resistant (TR-ACP) |  | 10.5% | 15.7%
| ACTH |  |  |  | +-2.0 up to 20 pmol/L | 10% >20 pmol/L
P- | Activated partial thromboplastine time |  | 4.5% | 6.7% |  |  | 20% | 15%
P- | Adiponectin |  | 29.1%
S- | Adenosine deaminase (ADA) |  | 16.7%
P- | Alanine |  | 26.6%
S- | Alanine aminotransferase (ALT) | 13.7% | 27.48% |  | 5 U/L ≤ 40 U/L | 12% ≥ 40 U/L | 20% | 20% | 15.1%
CSF- | Albumin, CSF |  |  |  | 0.02g/L≤0.45g/L | 5% ≥0.45g/L | 20%
S- | Albumin |  | 4.07% | 6.1% | 2g/L ≤ 33 g/L | 6% ≥ 33 g/L | 7% | 10% | 9%
U- | Albumin, concentration, first morning | 23.3% | 46.1% |  | 4.0 g/L ≤ 20 g/L | 20% ≥ 20 g/L | 15%
U- | Albumin, output, night urine |  | 40.6% |  | 4.0 g/L ≤ 20 g/L | 20% ≥ 20 g/L | 15%
S- | Albumin, glycated |  | 7.2% | 10.8%
S- | Alcohol/Ethanol |  |  |  | 2mmol/L ≤20mmol/L | 10% ≥20mmol/L |  | 25%
S- | Aldosterone |  | 36.7% |  | 24 pmol/L ≤ 160 pmol/L | 15% ≥ 160pmol/L
U- | Aldosterone | 19.8% | 46.56% |  |  |  | 25%
S- | Alkaline phosphatase (ALP) |  | 12.04% |  | 15 U/L ≤ 125 U/L | 12% ≥ 125 U/L | 18% | 30% | 18.5%
S- | Alkaline phosphatase, bone |  | 14.6% | 21.9%
S- | Alkaline phosphatase, liver |  | 15.4%
S- | Amikacin |  |  |  | 3.4umol/L ≤34umol/L | 10% ≥34umol/L | 20%
WB- | Ammonia |  |  |  | 5umol/L ≤50umol/L | 20% ≥50umol/L
U- | Ammonia, output, 24h |  | 29.6%
S- | Amyloid A |  | 37.1%
S- | Androstendione |  | 23.51% |  | 1.5 nmol/L ≤ 10 nmol/L | 15% ≥ 10nmol/L | 11.1%
S- | Angiotensin Converting Enzyme (ACE) |  |  |  | 1.0 U/L ≤ 10 U/L | 10% ≥ 10 U/L
S- | Anion gap |  | 11.3%
P- | Antithrombin III |  | 8.3% | 12.5% |  |  | 14%
S- | Apolipoprotein A1 |  | 9.1% | 13.6% |  |  | 12%
S- | Apolipoprotein B |  | 11.6% | 17.5% |  |  | 12%
P- | Arginine |  | 25.7%
S- | Arilestearase activity, non inhibited |  | 12.5%
P- | Ascorbate (Vitamin C) |  | 23.8%
S- | Ascorbate (Vitamin C) |  | 31.6%
S- | ASOT | [5 empty] | 25%
P- | Asparagine |  | 17.8%
S- | Aspartate aminotransferase (AST) | 8.3% | 16.69% |  | 5 U/L ≤ 40 U/L | 12% ≥ 40 U/L | 17% | 20% | 14.8%
P- | Aspartic acid |  | 41.6%
S- | β-2-Microglobulin |  | 9% | 13.5% | 0.2 mg/L ≤ 2.0 mg/L | 10% ≥ 2.0 mg/L | 12%
P- | β-Carotene |  | 27.7%
S- | β-Carotene |  | 43.1%
S- | β-Globulins |  | 11.7%
S- | β-hCG |  |  |  | 1.0 U/L ≤ 10 U/L | 10% ≥ 10 U/L | 4U/L or 20% | 3sd | 12.7%
B- | Base excess |  | 85%
S- | Basophile, count |  | 38.5%
S- | Bicarbonate |  |  |  | 2.0 umol/L ≤ 20 umol/L | 10% ≥ 20 umol/L | 3mmol/L or 15%
S- | Bilirubin total | 13.5% | 26.94% |  | 3.0 umol/L ≤ 25 umol/L | 12% ≥ 25 umol/L | 4umol/L or 20% | 6.8umol/L or 20% | 15.9%
S- | Bilirubin conjugated | 22.3% | 44.5% |  | 3.0 umol/L ≤ 15 umol/L | 20% ≥ 15 umol/L | 7umol/L or36% |  | 25.8%
P- | C Protein |  | 18.7% | 28%
S- | C reactive protein | 28.3% | 56.6% |  |  |  | 22%
S- | C reactive protein - high sensitivity |  | 66.54% |  |  |  | 22%
S- | C3 Complement |  | 8.4% |  |  |  | 12% | 3sd
S- | C4 Complement | 8% | 16% |  |  |  | 20% | 3sd
S- | CA 125 antigen | 17.7% | 35.4% |  | 6 kU/L ≤ 50kU/L | 12% ≥ 50kU/L | 20% |  | 12.9%
S- | CA 15.3 antigen | 10.4% | 20.8% |  | 3 kU/L ≤ 30kU/L | 10% ≥ 30kU/L | 20%
S- | CA 19.9 antigen | 23% | 46.03% |  | 6 kU/L ≤ 40kU/L | 15% ≥ 40kU/L | 20% |  | 16.5%
S- | CA 549 antigen |  | 16.2%
S- | Calcitonin | [5 empty] | 20%
S- | Calcium |  | 2.55% | 3.6% | 0.1 mmol/L ≤ 2.5 mmol/L | 4% ≥ 2.5 mmol/L | 4.4% | 0.25 mmol/L | 8.3%
S- | Calcium, complexed |  | 6.1%
U- | Calcium, concentration, 24h |  | 34.1% | 3.1% | 0.20mmol/L
B-Gas | Calcium, ionized |  |  |  | 0.05mmol/L |  | 10%
S- | Calcium, ionized |  | 2% |  |  |  | 4.4%
U- | Calcium, output, 24h | 15.5% | 31% |  |  |  | 12%
S- | Calcium, protein bound |  | 5.2%
S- | Calcium, ultrafiltrable |  | 2.7%
S- | Carbamazepine |  |  |  | 2.0 umol/L ≤ 20 umol/L | 10% ≥ 20 umol/L | 16% | 25%
S- | Carbohydrate deficient transferrin |  | 15.7% | 23.5%
B- | CO2, total |  | 4.86% | 7.3% |  |  | 9%
S- | Carcinoembryonic antigen (CEA) | 12.4% | 24.7% |  | 0.6 ug/L ≤ 5.0ug/L | 12% ≥ 5.0ug/L | 20% |  | 11.3%
S- | Carnitine, Acyl-free |  | 16.07%
S- | Carnitine, free |  | 11.26% | 16.9%
S- | Carnitine, total |  | 10.99%
S- | Ceruloplasmin (ferroxidase) |  | 7.9% | 11.9% |  |  | 12%
B-Gas | Chloride |  |  |  | 3mmol/L |  | 3.3%
CSF- | Chloride | [5 empty] | 20%
S- | Chloride |  | 1.5% | 2.2% | 3 mmol/L ≤ 100 mmol/L | 3% ≥ 100 mmol/L | 3.2% | 5% | 4.6%
U- | Chloride |  |  |  | 2 mmol/L ≤ 20 mmol/L | 10% ≥ 20 mmol/L | 10%
S- | Cholesterol | 4.5% | 9.01% |  | 0.3 mmol/L ≤ 5 mmol/L | 6% ≥ 5 mmol/L | 13% | 10% | 8.6%
S- | Cholinesterase, activity |  | 9.8% | 14.7% | 500 U/L ≤ 5000 U/L | 10% ≥ 5000 U/L | 20%
P- | Chromogranin A |  | 17.9%
P- | Citrulline |  | 29.9%
S- | Collagen type I C propeptide (PICP) |  | 13.4%
S- | Collagen type I N propeptide (PINP) |  | 20.5%
S- | Collagen type III N propeptide (PIIINP) |  | 33.3% | 49.9%
U- | Color, first morning |  | 39.6%
P- | Copper |  | 11.8%
S- | Copper |  | 7.47% | 11.2% |  |  | 13%
P- | Cortisol |  | 30.66%
S- | Cortisol | 11.4% | 22.8% |  | 30nmol/L ≤ 150nmol/L | 15% ≥ 150nmol/L | 25% | 25% | 13.5%
U- | Cortisol |  |  |  | 30nmol/L ≤ 200nmol/L | 15% ≥ 200nmol/L | 25%
S- | C Peptide | 10.4% | 20.8% |  | 0.15 up to 1.25nmol/L | 12% >1.25 nmol/L
S- | Creatine kinase (CK) | 28.3% | 30.3% |  |  | 12% ≥ 125 U/L | 20% | 30% | 12%
S- | Creatine kinase MB, % | 15.2% | 16.5% |  |  |  |  | 3sd
S- | Creatine kinase MB, activity | 8.3% | 24.1% |  | 3 U/L ≤ 15 U/L | 20% ≥ 15 U/L | 20% | 3sd
S- | Creatine kinase MB, mass | 15% | 30.06% |  | 3 ug/L ≤ 15ug/L | 20% ≥ 15 ug/L | 20% | 3sd
B-Gas | Creatinine |  |  |  | 10mmol/L ≤ 100mmol/L | 10% ≥100mmol/L
S- | Creatinine | 7.7% | 8.87% |  | 8 umol/L ≤ 100 umol/L | 8% ≥ 100 umol/L | 10% | 26umol/L or 15% | 12.6%
U- | Creatinine, concentration, 24h |  | 28.4% |  | 0.5 mmol/L ≤ 5 mmol/L | 10% ≥ 5 mmol/L | 15%
U- | Creatinine, concentration, first morning |  | 27.8% |  | 0.5 mmol/L ≤ 5 mmol/L | 10% ≥ 5 mmol/L | 15%
U- | Creatinine, concentration, random |  | 42.1% |  | 0.5 mmol/L ≤ 5 mmol/L | 10% ≥ 5 mmol/L | 15%
U- | Creatinine, output, 24h |  | 15.4% |  | 0.5 mmol/L ≤ 5 mmol/L | 10% ≥ 5 mmol/L | 15%
S- | C-Terminal telopeptide type I collagen (CTY I) |  | 15.45%
S- | Cyclosporin | [5 empty] | 25%
S- | Cyfra 21.1 Antigen | 13.9% | 27.9%
S- | Cystastin C |  | 7.61%
P- | Cysteine |  | 8.3% | 12.4%
P- | Cystine |  | 47%
S- | D-Dimer (MoM) |  | 28.04% |  |  |  | 20%
S- | Dehydroepiandrosterone sulfate (DHEAS) |  | 13.08% | 19.6% | 1.2 umol/L ≤ 10umol/L | 12% ≥10umol/L | 20% |  | 14.2%
U- | Deoxypyridinoline/creatinine, 24h |  | 21.9%
U- | Deoxypyridinoline/creatinine, first morning |  | 20.7%
U- | Deoxipyridinoline/minute, first morning |  | 21.2%
S- | Digoxin |  |  |  | 0.2 nmol/L ≤ 2.0 nmol/L | 10% ≥ 2.0 nmol/L | 18% | 20% or 0.26nmol/L
P- | Dipeptidyl-peptidase IV (ACE) |  | 10.9%
S- | Dipeptidyl-peptidase IV (ACE) |  | 17.9%
P- | Elastase |  | 15.11%
B- | Eosinophils, count |  | 37.1%
B- | Erythrocytes, count |  | 4.4% | 6.7% |  |  | 6% | 6%
B- | Erythrocyte distribution wide (RDW) |  | 4.6% |  |  |  | 5.6 or 4.6%
B- | Erythrocyte sedimentation rate (ESR) | [5 empty] | 30%
S- | Estradiol | 13.4% | 26.86% |  | 25 pmol/L ≤ 100 pmol/L | 25% ≥ 100.0 pmol/L | 20% |  | 13.7%
S- | Estradiol, free |  |  |  | 0.9 pmol/L ≤ 6.0 pmol/L | 15% ≥ 6.0 pmol/L
S- | Ethosuxamide | [5 empty] | 12%
P- | Factor II coagulation | [5 empty] | 15%
P- | Factor V coagulation | [5 empty] | 15%
P- | Factor VII coagulation |  | 10.7% | 16.1% |  |  | 15%
P- | Factor VIII coagulation |  | 8.9% | 13.3% |  |  | 15%
P- | Factor IX | [5 empty] | 15%
P- | Factor X coagulation | [5 empty] | 15%
P- | Factor XI | [5 empty] | 15%
P- | Factor XII | [5 empty] | 15%
S- | Ferritin | 8.4% | 16.9% |  | 4.0ug/L ≤ 27.0ug/L | 15% ≥ 27ug/L | 15% |  | 13.1%
P- | Fibrinogen |  | 13.6% |  |  |  | 25% | 20%
(B)Erthry- | Folate |  | 26.7%
S- | Folate | 19.5% | 39% |  | 1.5nmol/L ≤ 6.0nmol/L | 25% ≥ 6.0nmol/L | 25% |  | 23.9%
S- | Follicle stimulating hormone (FSH) | 10.6% | 21.19% |  | 1.0IU/L ≤ 10.0 IU/L | 10% ≥ 10IU/L | 20% |  | 11.4%
S- | Fructosamine |  | 4.5% | 6.8% | 15 umol/L ≤ 250 umol/L | 6% ≥ 250 umol/L | 7%
S- | Galactosyl hydroxylysine |  | 16.8% |  |  |  | 25%
P- | γ-Fibrinogen |  | 19.3%
S- | γ-Globulins |  | 16.8%
S- | γ-glutamyltransferase (GGT) | 11.1% | 22.11% |  | 5 U/L ≤ 40 U/L | 12% ≥ 40 U/L | 21% |  | 17.9%
S- | Gastrin |  |  |  | 10 pmol/L ≤ 100 pmol/L | 10% ≥ 100 pmol/L | 20%
S- | Gentamicin |  |  |  | 0.2mg/L ≤ 2mgl/L | 10% ≥ 5.3mg/L | 20% | 25%
S- | Globulins, total |  | 8% | 12.1%
B-Gas | Glucose |  |  |  | 1.0mmol/L ≤ 10mmol/L | 10% ≥10mmol/L | 7.4%
CSF- | Glucose |  |  |  | 0.2mmol/L ≤ 2.0mmol/L | 10% ≥2.0mmol/L | 20%
P- | Glucose |  | 5.5% |  | 0.4mmol/L ≤ 5.0mmol/L | 8% ≥ 5.0mmol/L | 7.4% | 0.3mmol/L or 10% | 8.5%
S- | Glucose |  | 6.96%
U- | Glucose |  |  |  | 1.0mmol/L ≤ 10mmol/L | 10% ≥10mmol/L | 10%
(B)Erythr- | Glucose-6-phosphate-1-dehydrogenase (G6PDH) |  | 38.5%
B - spot | Glucose-6-phosphate-1-dehydrogenase (G6PDH) |  | 9.2%
P- | Glutamic acid |  | 61.4%
P- | Glutamine |  | 16.3%
S- | Glutathion peroxidase |  | 11.7% | 17.5%
P- | Glycine |  | 20.2%
S- | Growth Hormone |  |  |  | +-1 up to ≤ 7 mU/L | 15% ≥ 7 mU/L
P- | Haptoglobin | 13.6% | 25.1%
S- | Haptoglobin |  | 27.3% |  |  |  | 12%
S- | HDL cholesterol |  | 11.63% |  | 0.1 mmol/L ≤ 0.8 mmol/L | 12% ≥ 0.8 mmol/L | 12% | 30% | 21.1%
S- | HDL 1 cholesterol |  | 11.5% | 17.2%
S- | HDL 2 cholesterol |  | 23.9%
S- | HDL 3 cholesterol |  | 9.8% | 14.6%
B- | Hematocrit |  | 3.97% | 5.9% |  |  | 6% | 6%
B- | Hemoglobin |  | 4.19% | 6.3% |  |  | 7% | 7%
B- | Hemoglobin A1 C (NSP) (%) |  | 3% | 4.5% | 0.5 ≤ 10% | 5% ≥ 10% | 20%
B- | Hemoglobin A1 C (IFCC) (mmol/mol) |  | 3% | 4.5% | 4 ≤ 86mmol/mol | 5% ≥ 86 mmol/mol | 20%
B- | Hemoglobin A2 |  | 2.51%
P- | Histidine |  | 15.2%
P- | Homocysteine |  | 15.48% | 26.6% | 1.5umol/L ≤ 15 umol/L | 10% ≥ 15 umol/L | 25%
P- | Hydroxyproline |  | 45.1%
U- | Hydroxyproline/minute, first morning |  | 43%
U- | Hydroxyproline/minute, second void |  | 46.5%
S- | Immunoglobulin A | 6.8% | 13.5% |  |  |  | 19% | 3sd
S- | Immunoglobulin E | [5 empty] | 12% | 3sd
CSF- | Immunoglobulin G |  |  |  | 0.02g/L ≤ 0.1g/L | 20% ≥0.1g/L | 20% | 20%
S- | Immunoglobulin G |  | 8% |  |  |  | 12% | 25%
S- | Immunoglobulin M | 8.4% | 16.8% |  |  |  | 22% | 3sd
S- | Immunoglobulins κ chains |  | 8% | 12%
S- | Immunoglobulins λ chains |  | 8.6% | 12.9%
S- | Inhibin A | [5 empty] | 20%
S- | Inhibin B |  | 14.98%
S- | Insulin |  | 32.9% |  | 0.6 mU/L ≤ 5.0mU/L | 12% ≥ 5.0 mU/L | 25% |  | 15.3
S- | Insulin-like growth factor (IGF-1) |  | 24% |  | 3 up to 25 nmol/L | 12% >25nmol/L
S- | Insulin-like growth factor binding protein 3 (IGFBP-3) |  | 24.5%
S- | Intercellular adhesion molecule-1 (ICAM-1) |  | 6.8%
(B)Leuc- | Interferon receptor |  | 17.7%
S- | Interleukin 1-β |  | 36.5%
S- | Interleukin-8 |  | 29.6%
S- | Iron | 15.3% | 30.7% |  | 3 umol/L ≤ 25 umol/L | 12% ≥ 25 umol/L | 20% | 20% | 10%
P- | Isoleucine |  | 24.8%
S- | Kallicrein 6 |  | 17.24%
B-Gas | Lactate |  |  |  | 1.0mmol/L ≤ 10mmol/L | 10% ≥10mmol/L | 20%
B- | Lactate | 15.2% | 30.4% |  | 0.5 mmol/L ≤ 4.0 mmol/L | 12% ≥ 4.0 mmol/L | 20%
CSF- | Lactate |  |  |  | 0.3mmol/L ≤3.0mmol/L | 10% ≥ 3.0mmol/L | 20%
S- | Lactate dehydrogenase (LDH) (L to P) | 5.7% | 11.4% |  | 20 U/L ≤ 250 U/L | 8% ≥ 250 U/L | 15% | 20% | 12.8%
S- | Lactate dehydrogenase (LDH) | 5.7% | 11.4% |  |  |  | 15% | 20%
S- | Lactate dehydrogenase 1 isoform (LDH1) |  | 4.1% | 6.1% |  |  |  | 30%
S- | Lactate dehydrogenase 2 isoform (LDH2) |  | 3.7% | 5.6% |  |  |  | 30%
S- | Lactate dehydrogenase 3 isoform (LDH3) |  | 3.5% | 5.2% |  |  |  | 30%
S- | Lactate dehydrogenase 4 isoform (LDH4) |  | 6.9% | 10.3% |  |  |  | 30%
S- | Lactate dehydrogenase 5 isoform (LDH5) |  | 9.7% | 14.6% |  |  |  | 30%
P- | Lactoferrin |  | 16.4%
S- | LDL Cholesterol |  | 11.9% | 17.8% | 0.2 mmol/L ≤ 2.0 mmol/L | 10% ≥ 2.0 mmol/L | 20%
P- | LDL Cholesterol (oxidized) |  | 30.9%
S- | LDL Cholesterol, small dense |  | 13%
S- | LDL receptor mRNA |  | 24.1%
P- | Leucine |  | 23.8%
B- | Leukocytes count |  | 15.49% |  |  |  | 15% | 15%
S- | Lipase |  | 37.88% |  | 12 U/L ≤ 60 U/L | 20% ≥ 60 U/L | 15% |  | 25.7%
S- | Lipoprotein (a) |  | 24.1% |  |  |  | 21%
S- | Lithium |  |  |  | 0.2 mmol/L |  | 0.2mmol/L or 12% | 20% or 0.3mmol/L
P- | Lutein |  | 16.9%
S- | Luteinizing hormone (LH) | 14% | 27.92% |  | 1.5 IU/L ≤ 10.0 IU/L | 15% ≥ 10.0 IU/L | 20% |  | 12.1%
P- | Lycopene |  | 28.1%
B- | Lymphocytes, count |  | 17.6%
P- | Lysine |  | 19.5%
(B)Erythr- | Magnesium |  | 7.8% | 11.7%
(B)Leuc- | Magnesium |  | 20.83%
(B)Mon - | Magnesium |  | 21.7%
S- | Magnesium |  | 4.8% | 7.2% | 0.1 mmol/L ≤ 1.25 mmol/L | 8% ≥ 1.25 mmol/L | 8% | 25% | 11%
U- | Magnesium, ionized |  | 2.9%
U- | Magnesium, output, 24h |  | 45% |  | 0.2 mmolL |  | 25%
(B)Erythr- | Mean corpuscular hemoglobin (MCH) |  | 2.5% | 3.8% |  |  | 6%
(B)Erythr- | Mean corpuscular hemoglobin concentration (MCHC) |  | 1.27% | 1.9% |  |  | 6%
(B)Erythr- | Mean corpuscular volume (MCV) |  | 2.42% | 2.3% |  |  | 6%
(B)Plat- | Mean platelet volume (MPV) |  | 5.84% | 8.8% |  |  | 20%
S- | Methotrexate | [5 empty] | 20%
P- | Metionine |  | 23.6%
B- | Monocytes, count |  | 27.9%
S- | Myeloperoxidase |  | 41.4%
S- | Myoglobin | 13.5% | 19.6% |  |  |  | 28%
U- | N-Acetyl Glucosaminidase, concentration, first morning |  | 58%
U- | N-Acetyl Glucosaminidase/Creatinine |  | 56%
B- | Neutrophyles, count |  | 23.35%
U- | Nitrogen, output |  | 18.4%
U- | N-Telopeptide type I collagen |  | 22.95%
S- | N-terminal (NT)-proBNP |  | 13% |  | 25ng/L≤125ng/L | 20%≥125ng/L | 20%
P- | Ornithine |  | 29.7%
P- | Osmolality |  | 1.6%
Saliva- | Osmolality |  | 17.1%
S- | Osmolality |  | 1.5% | 2.3% | 8 mmol/kg ≤ 266 mmol/kg | 3% ≥ 266 mmol/kg | 3%
U- | Osmolality, first morning | 19.7% | 39.5% |  | 6 mmol/kg ≤ 300 mmol/kg | 2% ≥ 300 mmol/kg | 5%
S- | Osteocalcin |  | 13.13% | 19.7%
U- | Oxalate, concentration, 24h |  | 48.2% |  |  |  | 25%
U- | Oxalate, output, 24h |  | 46.8% |  |  |  | 25%
B-Gas | pCO2 |  | 5.7% |  | 2.0mmHg ≤ 25mmHg | 8% ≥ 25mmHg | 8% | 0.7kPa or 8%
B- Gas | pH [H+] | 1.9% | 3.9% |  | 0.04 units |  | 0.5% | 0.04 units
B- Gas | pO2 |  |  |  | 5.0mmHg ≤ 100mmHg | 5% ≥ 100mmHg | 12% | 3sd
S- | Paraoxonase 1 |  | 32.3%
S- | Paraoxonase 1 substrate inhibition (PON 4SI) |  | 23.2%
S- | Paraoxonase, activity (salt stimulated) |  | 28.3%
P- | Parathyroid hormone (PTH) |  | 33.43% | [5 empty] | 21.2%
S- | Parathyroid hormone (PTH) | 15.1% | 30.2% |  | 1.0 pmol/L ≤ 8.0 pmol/L | 12% ≥ 8.0 pmol/L | 25%
S- | Phenobarbital |  |  |  | 3 umol/L ≤ 30 umol/L | 10% ≥ 30 umol/L | 16% | 20%
S- | Phenylacetate |  | 12%
P- | Phenylalanine |  | 18.3%
S- | Phenytoin |  |  |  | 3 umol/L ≤ 30 umol/L | 10% ≥ 30 umol/L | 16% | 25%
S- | Phosphate | 5.1% | 10.11% |  | 0.06 mmol/L ≤ 0.75 mmol/L | 8% ≥ 0.75 mmol/L | 13% |  | 9.4%
U- | Phosphate, output, 24h | 11% | 22.1% |  | 2.5mmol/L |  | 13%
Patient- | Phosphate tubular reabsorption |  | 3.3%
S- | Phospholipids |  | 8.6%
P- | Plasminogen | [5 empty] | 15%
B- | Platelets, count |  | 13.4% |  |  |  | 19% | 25%
B- | Platelet distribution wide
B- | Plateletcrit | [5 empty] | 20%
U- | Porphobilinogen |  | 22.9%
U- | Porphyrins (total)
B-Gas | Potassium |  |  |  | 0.2mmol/L |  | 8%
(B)Leuc- | Potassium |  | 16%
S- | Potassium |  | 5.61% | 8.4% | 0.2 mmol/L ≤ 4.0 mmol/L | 5% ≥ 4.0 mmol/L | 8% | 0.5mmol/L | 5.6%
U- | Potassium, output | 14.2% | 28.4% |  | 2 mmol/L ≤ 20.0 mmol/L | 10% ≥ 20.0 mmol/L | 10%
S- | Prealbumin |  | 14.5% |  |  |  | 12%
S- | Pregnancy-associated plasma protein A (PAPP-A) |  | 15.1% |  |  |  | 20%
S- | Primidone | [5 empty] | 20%
S- | Procalcitonin | [5 empty] | 20%
S- | Progesterone |  |  |  | 2.0 nmol/L ≤ 10.0 nmol/L | 15% ≥ 10.0 nmol/L | 25% | 25% | 15.6%
P- | Prolactin | 14.7% | 51.34%
S- | Prolactin |  | 29.4% |  | 40 mIU/L ≤ 400 mIU/L | 10% ≥ 400 mIU/L | 20% |  | 17.1%
P- | Proline |  | 40.5%
P- | Prolyl endopeptidase |  | 19.3%
S- | Properdin factor B |  | 11.5%
S- | Prostatic specific antigen (PSA) Free |  |  |  | 0.2 ug/L ≤ 1.4 ug/L | 15% ≥ 1.4 ug/L | 25% |  | 12.2%
S- | Prostatic specific antigen (PSA) Total | 16.8% | 33.6% |  | 0.4 ug/L ≤ 5.0 ug/L | 8% ≥ 5.0 ug/L | 15% |  | 11.6%
CSF- | Protein |  |  |  | 0.02g/L ≤ 0.45g/L | 5% ≥ 0.45g/L | 20%
S- | Protein |  | 3.63% | 5.4% | 3 g/L ≤ 60 g/L | 5% ≥ 60 g/L | 6% | 10% | 8.7%
S- | Protein, glycated |  | 3.7% | 5.5%
U- | Protein, output, 24h | 20% | 40% |  | 0.1 g/L ≤ 1.0 g/L | 10% ≥ 1.0 g/L | 0.1g/L or 10%
P- | Protein C activity | [5 empty] | 15%
P- | Protein S activity | [5 empty] | 15%
P- | Prothrombin time |  | 5.3% | 7.9% |  |  | 20% | 15%
U- | Pyridinoline |  | 23.6%
B- | Pyruvate |  | 17.5%
Patient- | Reabsorption tubular phosphate |  |  | 4.9%
S- | Receptor for advanced glycation end-products (RAGE) |  | 26.63%
B- | Red cell distribution wide (RDW) |  | 4.6% | 6.8%
P- | Renin Activity | [5 empty] | 20%
P- | Renin concentration | [5 empty] | 20%
S- | Reticulocyte highly fluorescent, count |  | 24%
S- | Reticulocyte low fluorescent, count |  | 2.6%
S- | Reticulocyte medium fluorescent, count |  | 19.6%
S- | Reticulocyte, count |  | 16.8%
P- | Retinol |  | 10.6%
S- | Retinol |  | 17.1% |  |  |  | 25%
S- | Rheumatoid factor |  | 13.5% |  |  |  | 24%
S- | SCC antigen |  | 45.8%
P- | S Protein |  | 20.7% | 31.1%
S- | Salicylate |  |  |  | 0.1 mmol/L ≤ 1.0 mmol/L | 10% ≥ 1.0 mmol/L | 12%
P- | Selenium |  | 14.5%
B- | Selenium |  | 14.1%
P- | Serine |  | 21.7%
S- | Sex hormone binding globulin (SHBG) |  | 20.42% |  | 6.0 nmol/L ≤ 50.0 nmol/L | 12% ≥ 50.0 nmol/L | 21% |  | 13.7%
(B)Erythr- | Sodium |  | 4.6% | 6.9%
B-Gas | Sodium |  |  |  | 3.0mmol/L
(B)Leuc- | Sodium |  | 57.7%
CSF- | Sodium | [5 empty] | 20%
S- | Sodium |  | 0.73% | 1.1% | 3 mmol/L ≤ 150 mmol/L | 2% ≥ 150 mmol/L | 3% | 4 mmol/L | 3.6%
B- | Sodium Bicarbonate |  | 4.9%
S- | Sodium Bicarbonate |  | 5.6%
Sweat- | Sodium Chloride |  | 19.7%
U- | Sodium, output, 24 h. | 16% | 32% |  | 2 mmol/L ≤ 20 mmol/L | 10% ≥ 20 mmol/L | 10%
P- | Soluble CD163 |  | 16.7%
U- | Specific gravity |  | 0.6%
Semen- | Spermatozoa, concentration |  | 37.7%
Semen- | Spermatozoa, morphology |  | 28.2%
Semen- | Spermatozoa, progressive motility |  | 21.6%
Semen- | Spermatozoa, fast progressive motility |  | 29.3%
Semen- | Spermatozoa, total motility |  | 23.9%
Semen- | Spermatozoa, vitality |  | 15.4%
S- | Superoxide dismutase |  | 19.1%
(B)Erythr- | Superoxide dismutase |  | 13.5%
P- | Taurine |  | 38.6%
P- | Testosterone |  | 21.07%
S- | Testosterone |  | 13.61% |  | 0.4 nmol/L ≤ 2.7 nmol/L | 15% ≥ 2.7 nmol/L | 22% |  | 15.5%
Saliva- | Testosterone |  | 22.7%
S- | Theophylline |  |  |  | 3 umol/L ≤ 30 umol/L | 10% ≥ 30 umol/L | 20% | 25%
S- | Thyroglobulin | 11% | 21.9% |  | 0.2 ug/L ≤ 1.7 ug/L | 12% ≥ 1.7 ug/L | 17.8% |  | 18.1%
S- | Thyroglobulin antibody |  | 27.6%
S- | Thyroid peroxidase antibody |  | 46.2%
P- | Thyroid stimulating hormone (TSH) |  | 38.2%
S- | Thyroid stimulating hormone (TSH) | 11.9% | 23.7% |  | 0.1 mU/L ≤ 0.5mU/L | 20% ≥ 0.5 mU/L | 20% | 3sd | 11.2%
P- | Thrombin Time | [5 empty] | 15%
S- | Thyroxine (T4) |  | 7% | 10.4% | 12 nmol/L ≤ 120 nmol/L | 10% ≥ 120 nmol/L | 15% | 20%
P- | Thyroxine, free (FT4) |  | 8.74%
S- | Thyroxine, free (FT4) |  | 8% |  | 1.5 pmol/L ≤ 12 pmol/L | 12% ≥ 12 pmol/L | 16% | 3sd | 11.9%
S- | Thyroxine binding globulin (TBG) |  | 0.1% | 0.2%
P- | Tirosine |  | 24.1%
S- | Tissue polypeptide antigen (TPA) |  | 43.4%
S- | Tissue polypeptide specific antigen (TPS) |  | 58.3%
S- | Tobramycin |  |  |  | 0.5mg/L ≤5mgl/L | 10% ≥5.1mg/L | 20% | 25%
U- | Total catecolamines, concentration, 24h |  | 29.8%
S- | Transferrin |  | 3.8% | 5.7% | 0.2 g/L ≤ 2.5 g/L | 8% >2.50 g/L | 11%
P- | Treonine |  | 24.2%
S- | Triglyceride | 13% | 25.99% |  | 0.2 mmol/L ≤ 1.6 mmol/L | 12% ≥ 1.6 mmol/L | 20% | 25% | 15.4%
P- | Triiodothyronine (T3) |  | 12.94%
S- | Triiodothyronine (T3) |  | 9.22% |  | 0.2 nmol/L ≤ 1.3 nmol/L | 15% ≥ 1.3 nmol/L | 16%
S- | Triiodothyronine, free (FT3) |  | 11.3% |  | 0.7 pmol/L ≤ 3.5 pmol/L | 20% ≥ 3.5 pmol/L | 20% |  | 13.3%
P- | Troponin I |  | 76.36%
S- | Troponin I |  | 27.91% |  | 0.002 ug/mL ≤ 0.01 ug/mL | 20% ≥ 0.010 ug/mL | 23%
S- | Troponin T |  | 48.9% |  | 0.01 ug/L ≤ 0.05 ug/L | 20% ≥ 0.05 ug/L | 23%
P- | Tryptophan |  | 57.3%
S- | Tumor Necrosis Factor-a (TNF-a) |  | 48.4%
S- | Urate | Fso | 11.97% |  | 0.03 mmol/L ≤ 0.38 mmol/L | 8% ≥ 0.38 mmol/L | 15% | 17% | 10.6%
U- | Urate, output, 24h | 9.7% | 19.39% |  | 0.3mmol/L |  | 15%
B-Gas | Urea |  |  |  | 1.0mmol/L ≤10mmol/L | 10% ≥10mmol/L
S- | Urea | 7.8% | 15.55% |  | 0.5 mmol/L ≤ 4.0 mmol/L | 12% ≥ 4.0 mmol/L | 15% | 0.7mmol/L or 9% | 12.5%
U- | Urea, output, 24h | 11% | 22.1% |  | 20 mmol/L ≤ 200 mmol/L | 10% ≥ 200 mmol/L | 16%
P- | Valine |  | 19.1%
S- | Valproate |  |  |  | 25 umol/L ≤ 250 umol/L | 10% ≥ 250 umol/L | 17% | 25%
S- | Vancomycin |  |  |  | 2.0mg/L ≤20.3mgl/L | 10% ≥20.3mg/L | 20%
U- | Vanilmandelic Acid concentration, 24h |  | 31.3% |  |  |  | 25%
S- | Vascular cell adhesion molecule-1 (VCAM-1) |  | 8.5% | 12.7%
P- | Vascular endotelial growth factor |  | 17.4% | 26.1%
B- | Vascular endotelial growth factor |  | 19.8%
S- | Vascular endotelial growth factor |  | 21%
P- | Vitamin B1 |  | 7.2%
B- | Vitamin B2 (Riboflavin) |  | 7.7%
(B)Eryth- | Vitamin B2 (Riboflavin) |  | 8.5%
(B)Eryth- | Vitamin B2 status (gluthation reductase activation) |  | 14.4%
(B)Eryth- | Vitamin B12 |  | 30%
S- | Vitamin B12 |  |  |  | 18 pmol/L ≤ 120 pmol/L | 15% ≥ 120 pmol/L | 20% |  | 17.3%
(B)Eryth- | Vitamin B6 |  | 18.5%
B- | Vitamin B6 |  | 26.4%
(B)Eryth- | Vitamin E (Tocopherol) |  | 11.9%
(B)Eryth- | Vitamin K (Phylloquinone) |  | 45.9%
P- | Von Willebrand factor |  | 8.9% | 13.4%
S- | Water |  | 3.3% | 5%
S- | Zinc |  | 11% |  |  |  | 14%
P- | Zinc |  | 13.5%

## Reference values
Algorithm: | A: Westgard's recommended choice | 61
| B: Published target: NGSP or ClinChem | 2
| C: Algorithm: 1. SOTA 2. EFLM_M 3. BV_D 4. RCPA 5. Rilibak | 35
SEQ | SIGMA SHEET: ANALYTE | WESTGARD | Recomm. Choice | CLIA | CLIA '19 | RIQAS SOTA | RICOS | Spanish Min | EFLM (Des:f; Min:g) | . | EFLM | Matrix | Within-subject median | Between-subject median | Date Updated on EFLM | Date Scraped | Minimum CV | Minimum Bias | Minimum TEa | Desirable CV | Desirable Bias | Desirable TEa | Optimum CV | Optimum Bias | Optimum TEa | Biological Variation | Matrix | OptimalTEa | Desirable TEa | Minimum TEa | RCPA Lower Goal | RCPA Upper Goal | Rilibak | New Target | Choice | Source
1 | [9 empty] | . | 17-a-hydroxyprogesterone | Serum/plasma | 28.3 | 38.5 | June 6th 2022 | 2022-06-08 | 21.225 | 17.9183 | 52.9396 | 14.15 | 11.9456 | 35.2931 | 7.075 | 5.9728 | 17.6465
2 | [9 empty] | . | 25-hydroxy vitamin D3 | Serum/plasma | 6.8 | 26.1 | June 6th 2022 | 2022-06-08 | 5.1 | 10.1142 | 18.5292 | 3.4 | 6.7428 | 12.3528 | 1.7 | 3.3714 | 6.1764
3 | ALPHA 1-ANTITRYPSIN | Alpha-1 Antitrypsin | ± 9.6% |  |  | ±9.6% | ± 9.2% |  | ± 9.30% [g] | . | a-1 antitrypsin | Serum/plasma | 4.1 | 10.5 | June 6th 2022 | 2022-06-08 | 3.075 | 4.227 | 9.3008 | 2.05 | 2.818 | 6.2005 | 1.025 | 1.409 | 3.1003 | α1-Antitrypsin | S- |  | 9.2% | 13.8% |  |  |  | 9.6 | A: Westgard's recommended choice | SOTA
4 | [9 empty] | . | Actual bicarbonate | Whole Blood | 4 | 4.8 | February 11th 2021 | 2022-06-08 | 3 | 2.3431 | 7.2931 | 2 | 1.562 | 4.862 | 1 | 0.78102 | 2.431
5 | [9 empty] | . | Adiponectin | Serum/plasma | 18.8 | 51.2 | October 20th 2019 | 2022-06-08 | 14.1 | 20.4534 | 43.7184 | 9.4 | 13.6356 | 29.1456 | 4.7 | 6.8178 | 14.5728
6 | [9 empty] | . | Adrenaline | No Matrix Set | 135 | 54 | February 14th 2020 | 2022-06-08 | 101.25 | 54.5248 | 221.5873 | 67.5 | 36.3499 | 147.7249 | 33.75 | 18.1749 | 73.8624
7 | ALANINE TRANSAMINASE | Alanine aminotransferase (ALT) | ±16.1% | ± 20% | ± 15% | ±15.1% | ± 27.5% | ± 23% | ±14.4%[a] ±16.1%[f] ±24.1%[g] | . | Alanine transaminase (ALT) | Serum/plasma | 10 | 29.3 | June 6th 2022 | 2022-06-08 | 7.5 | 11.6098 | 23.9848 | 5 | 7.7399 | 15.9899 | 2.5 | 3.8699 | 7.9949 | Alanine aminotransferase (ALT) | S- | 13.7% | 27.48% |  | 5 U/L ≤ 40 U/L | 12% ≥ 40 U/L |  | 16.1 | A: Westgard's recommended choice | EFLM_D
8 | ALBUMIN | Albumin | ± 10% | ± 10% | unchanged | ±9% | ± 4.1% | ± 14% | ±5.2%[g] | . | Albumin | Serum/plasma | 2.5 | 4.9 | June 6th 2022 | 2022-06-08 | 1.875 | 2.0628 | 5.1566 | 1.25 | 1.3752 | 3.4377 | 0.625 | 0.68761 | 1.7189 | Albumin | S- |  | 4.07% | 6.1% | 2g/L ≤ 33 g/L | 6% ≥ 33 g/L |  | 10 | A: Westgard's recommended choice | CLIA
9 | ALDO | Aldosterone | [6 empty] | ±42.8%[f] ±64.2%[i] | . | Aldosterone | Serum/plasma | 36.6 | 34.7 | October 20th 2019 | 2022-06-08 | 27.45 | 18.913 | 64.2055 | 18.3 | 12.6087 | 42.8037 | 9.15 | 6.3043 | 21.4018 | Aldosterone | S- |  | 36.7% |  | 24 pmol/L ≤ 160 pmol/L | 15% ≥ 160pmol/L |  | 64.2 | C: Algorithm - choice 2 | EFLM_M
10 | [9 empty] | . | Alkaline phosphatase (ALP) bone isoenzyme | Serum/plasma | 6.6 | 35.6 | November 9th 2019 | 2022-06-08 | 4.95 | 13.5775 | 21.745 | 3.3 | 9.0517 | 14.4967 | 1.65 | 4.5258 | 7.2483
11 | [9 empty] | . | Alkaline phosphatase (ALP), liver type | Serum/plasma | 5.3 | 24 | June 6th 2022 | 2022-06-08 | 3.975 | 9.2168 | 15.7756 | 2.65 | 6.1446 | 10.5171 | 1.325 | 3.0723 | 5.2585
12 | ALPHA-FETO PROTEIN | Alpha- Fetoprotein (AFP) | ± 21.9% |  | ± 15% | ±13.3% | ± 21.9% | ± 20 | ± 34.83% ±17.5%[g] ±26.5%[i] | . | Alphafetoprotein | Serum/plasma | 4.5 | 55.1 | April 19th 2022 | 2022-06-08 | 3.375 | 20.7313 | 26.3 | 2.25 | 13.8209 | 17.5334 | 1.125 | 6.9104 | 8.7667 | α-Fetoprotein (AFP) | S- | 10.9% | 21.9% |  | 2 kIU/L ≤ 17 kIU/L | 12% ≥ 17kIU/L |  | 21.9 | A: Westgard's recommended choice | RICOS
13 | [5 empty] | ±20.3% | ± 14.6% | ± 35% |  | . | Amylase | Serum/plasma | 6.5 | 30.2 | June 6th 2022 | 2022-06-08 | 4.875 | 11.5843 | 19.6281 | 3.25 | 7.7229 | 13.0854 | 1.625 | 3.8614 | 6.5427
14 | [9 empty] | . | Antimullerian hormone | Serum/plasma | 19.2 |  | December 14th 2020 | 2022-06-08 | 14.4 | 7.2 | 30.96 | 9.6 | 4.8 | 20.64 | 4.8 | 2.4 | 10.32
15 | APOPROTEIN A | Apolipoprotein A1 | ± 11.3% |  |  |  | ± 9.1% |  | ± 7.56% [f] ± 11.3% [g] | . | Apolipoprotein A1 | Serum/plasma | 5.3 | 11.1 | June 6th 2022 | 2022-06-08 | 3.975 | 4.6127 | 11.1714 | 2.65 | 3.0751 | 7.4476 | 1.325 | 1.5376 | 3.7238 | Apolipoprotein A1 | S- |  | 9.1% | 13.6% |  |  |  | 11.3 | A: Westgard's recommended choice | EFLM_M
16 | APOPROTEIN B | Apolipoprotein B | ± 11.6% |  |  |  | ± 11.6% |  | ± 11.48% [f] ± 17.2% [g] | . | Apolipoprotein B | Serum/plasma | 7.3 | 20.1 | June 6th 2022 | 2022-06-08 | 5.475 | 8.0192 | 17.053 | 3.65 | 5.3461 | 11.3686 | 1.825 | 2.6731 | 5.6843 | Apolipoprotein B | S- |  | 11.6% | 17.5% |  |  |  | 11.6 | A: Westgard's recommended choice | RICOS
17 | ASPARTATE TRANSAMINASE | Aspartate aminotransferase (AST) | ± 16.7% | ± 20% | ± 15% | ±14.8% | ± 16.7% | ± 21% | ±13.4%[a] ±13.65%[f] ±20.5% [g] | . | Aspartate transaminase (AST) | Serum/plasma | 9.5 | 20.7 | June 6th 2022 | 2022-06-08 | 7.125 | 8.5409 | 20.2972 | 4.75 | 5.694 | 13.5315 | 2.375 | 2.847 | 6.7657 | Aspartate aminotransferase (AST) | S- | 8.3% | 16.69% |  | 5 U/L ≤ 40 U/L | 12% ≥ 40 U/L |  | 16.7 | A: Westgard's recommended choice | RICOS
18 | [9 empty] | . | b-2 microglobulin | Serum/plasma | 4.1 | 11.5 | June 6th 2022 | 2022-06-08 | 3.075 | 4.5784 | 9.6521 | 2.05 | 3.0523 | 6.4348 | 1.025 | 1.5261 | 3.2174
19 | [9 empty] | . | Basophils | Whole Blood | 12.4 | 26.3 | June 6th 2022 | 2022-06-08 | 9.3 | 10.9037 | 26.2487 | 6.2 | 7.2692 | 17.4992 | 3.1 | 3.6346 | 8.7496
20 | TOTAL BILIRUBIN | Bilirubin, total | ± 0.4 mg/dL or ± 20% | ± 0.4 mg/dL or ± 20% | ± 20% | ±15.9% | ± 27% | ± 24% | ±24.84%[b] ±37.3%[g] | . | Bilirubin | Serum/plasma | 20 | 26.6 | June 6th 2022 | 2022-06-08 | 15 | 12.48 | 37.23 | 10 | 8.32 | 24.82 | 5 | 4.16 | 12.41 | Bilirubin total | S- | 13.5% | 26.94% |  | 3.0 umol/L ≤ 25 umol/L | 12% ≥ 25 umol/L |  | 20 | A: Westgard's recommended choice | CLIA
21 | [9 empty] | . | Calcitonin | Serum/plasma | 13 | 65.8 | January 14th 2022 | 2022-06-08 | 9.75 | 25.152 | 41.2395 | 6.5 | 16.768 | 27.493 | 3.25 | 8.384 | 13.7465
22 | CALCIUM | Calcium, total | ± 1.0 mg/dL | ± 1.0 mg/dL | unchanged | ±8.3% | ± 2.5% |  | ± 2.3%[f] ± 3.4%[g] | . | Calcium (Ca) | Serum/plasma | 1.8 | 2.7 | June 6th 2022 | 2022-06-08 | 1.35 | 1.2169 | 3.4444 | 0.9 | 0.81125 | 2.2962 | 0.45 | 0.40562 | 1.1481 | Calcium | S- |  | 2.55% | 3.6% | 0.1 mmol/L ≤ 2.5 mmol/L | 4% ≥ 2.5 mmol/L |  | 8.3 | C: Algorithm - choice 1 | SOTA
23 | CA 125 | CA 125 | ± 35.4% |  | ± 20% | ±12.9% | ± 35.4% |  | ± 13.9%[f] ± 16.2%[g] ± 20.9%[i] | . | Cancer antigen 125 (CA-125) | Serum/plasma | 8.7 | 25.4 | June 6th 2022 | 2022-06-08 | 6.525 | 10.0682 | 20.8345 | 4.35 | 6.7122 | 13.8897 | 2.175 | 3.3561 | 6.9448 | CA 125 antigen | S- | 17.7% | 35.4% |  | 6 kU/L ≤ 50kU/L | 12% ≥ 50kU/L |  | 35.4 | A: Westgard's recommended choice | RICOS
24 | CA 19-9 | CA 19-9 | ± 46.0% |  |  | ±16.5% | ± 46.0% |  | ± 17.9%[f] ± 18.0%[g] ± 26.9%[i] | . | Cancer antigen 19-9 (CA 19-9) | Serum/plasma | 4.3 | 57.4 | June 6th 2022 | 2022-06-08 | 3.225 | 21.5853 | 26.9066 | 2.15 | 14.3902 | 17.9377 | 1.075 | 7.1951 | 8.9689 | CA 19.9 antigen | S- | 23% | 46.03% |  | 6 kU/L ≤ 40kU/L | 15% ≥ 40kU/L |  | 46 | A: Westgard's recommended choice | RICOS
25 | [9 empty] | . | Cancer antigen 72-4 (CA 72-4) | Serum/plasma | 50.3 | 103.4 | November 20th 2019 | 2022-06-08 | 37.725 | 43.1195 | 105.3658 | 25.15 | 28.7464 | 70.2439 | 12.575 | 14.3732 | 35.1219
26 | [9 empty] | . | Carbon dioxide (total) | Whole Blood | 4 | 4.8 | February 11th 2021 | 2022-06-08 | 3 | 2.3431 | 7.2931 | 2 | 1.562 | 4.862 | 1 | 0.78102 | 2.431 | CO2, total | B-
27 | CARCINOEMBRYONIC AG | Carcinogenic Embryonic Antigen (CEA) | ± 24.7% |  | ± 15.0% | ±11.3% | ± 24.7% | ± 16.0% | ± 26.94% ± 20.5%[f] ± 30.8%[i] | . | Carcinoembryonic antigen (CEA) | Serum/plasma | 6.8 | 59.3 | June 6th 2022 | 2022-06-08 | 5.1 | 22.3832 | 30.7982 | 3.4 | 14.9222 | 20.5322 | 1.7 | 7.4611 | 10.2661 | Carcinoembryonic antigen (CEA) | S- | 12.4% | 24.7% |  | 0.6 ug/L ≤ 5.0ug/L | 12% ≥ 5.0ug/L |  | 24.7 | A: Westgard's recommended choice | RICOS
28 | [9 empty] | . | Ceruloplasmin | Serum/plasma | 4.9 | 15.2 | May 26th 2020 | 2022-06-08 | 3.675 | 5.9889 | 12.0526 | 2.45 | 3.9926 | 8.0351 | 1.225 | 1.9963 | 4.0175
29 | CHLORIDE | Chloride | ± 5% | ± 5% | unchanged | ±4.6% | ± 1.5% | ± 9% | ± 1.2% [b] ± 1.3% [f] ± 2.0% [g] | . | Chloride | Serum/plasma | 1 | 1.3 | June 6th 2022 | 2022-06-08 | 0.75 | 0.61505 | 1.8525 | 0.5 | 0.41003 | 1.235 | 0.25 | 0.20502 | 0.61752 | Chloride | S- |  | 1.5% | 2.2% | 3 mmol/L ≤ 100 mmol/L | 3% ≥ 100 mmol/L |  | 5 | A: Westgard's recommended choice | CLIA
30 | TOTAL CHOLESTEROL | Cholesterol, total | ± 10% | ± 10% | unchanged | ±8.6% | ± 9.0% | ± 11% | ± 9.3%[d] ± 8.7%[f] ±13.0%[g] | . | Cholesterol | Serum/plasma | 5.3 | 16.2 | June 6th 2022 | 2022-06-08 | 3.975 | 6.3919 | 12.9506 | 2.65 | 4.2612 | 8.6337 | 1.325 | 2.1306 | 4.3169 | Cholesterol | S- | 4.5% | 9.01% |  | 0.3 mmol/L ≤ 5 mmol/L | 6% ≥ 5 mmol/L |  | 10 | A: Westgard's recommended choice | CLIA
31 | [9 empty] | . | Chromogranin A | Serum/plasma | 14.8 | 30.6 | June 6th 2022 | 2022-06-08 | 11.1 | 12.7467 | 31.0617 | 7.4 | 8.4978 | 20.7078 | 3.7 | 4.2489 | 10.3539
32 | COMPLEMENT C3 | Complement C3 | ± 15% |  | ± 15% | ±7.9% | ± 8.4% |  | ± 7.77%[f] ± 11.6%[g] | . | Complement 3 (C3) | Serum/plasma | 4.6 | 15.2 | November 25th 2019 | 2022-06-08 | 3.45 | 5.9553 | 11.6478 | 2.3 | 3.9702 | 7.7652 | 1.15 | 1.9851 | 3.8826 | C3 Complement | S- |  | 8.4% |  |  |  |  | 15 | A: Westgard's recommended choice | CLIA '19
33 | COMPLEMENT C4 | Complement C4 | ± 16% | ± 5 mg/dL or 30% |  | ± 9% | ± 16% |  | ± 12.1%[f] ± 18.1%[g] | . | Complement 4 (C4) | Serum/plasma | 6.9 | 24.5 | April 29th 2020 | 2022-06-08 | 5.175 | 9.5449 | 18.0837 | 3.45 | 6.3633 | 12.0558 | 1.725 | 3.1816 | 6.0279 | C4 Complement | S- | 8% | 16% |  |  |  |  | 16 | A: Westgard's recommended choice | RICOS
34 | [9 empty] | . | Copper | Serum/plasma | 7.4 | 13.2 | June 6th 2022 | 2022-06-08 | 5.55 | 5.6748 | 14.8323 | 3.7 | 3.7832 | 9.8882 | 1.85 | 1.8916 | 4.9441
35 | CORTISOL II | Cortisol_I | ± 22.8% | ± 25.0% | ± 20.0% | ±13.5% | ± 22.8% | ± 28.0% | ± 26.3%[f] ± 39.3%[i] | . | Cortisol | Serum/plasma | 16.3 | 48.7 | June 6th 2022 | 2022-06-08 | 12.225 | 19.2583 | 39.4295 | 8.15 | 12.8389 | 26.2864 | 4.075 | 6.4194 | 13.1432 | Cortisol | S- | 11.4% | 22.8% |  | 30nmol/L ≤ 150nmol/L | 15% ≥ 150nmol/L |  | 22.8 | A: Westgard's recommended choice | RICOS
36 | C-REACTIVE PROTEIN | C-Reactive Protein (CRP) | ± 50.7% ± 72.9% (hs) |  | ± 1 mg/dL or 30% (hs) | ±8.8% | ± 56.6% |  | ± 50.7%[f] ± 72.9% hs [f] ± 76.1%[g] ± 109.4% hs [g] | . | C-reactive protein (CRP) | Serum/plasma | 34 | 83.6 | June 6th 2022 | 2022-06-08 | 25.5 | 33.8435 | 75.9185 | 17 | 22.5624 | 50.6124 | 8.5 | 11.2812 | 25.3062 | C reactive protein | S- |  | 56.6% |  |  |  |  | 50.7 | A: Westgard's recommended choice | EFLM_D
37 | [9 empty] | . | C-reactive protein (CRP) - high-sensitive | Serum/plasma | 58.9 | 77.4 | January 12th 2020 | 2022-06-08 | 44.175 | 36.4734 | 109.3621 | 29.45 | 24.3156 | 72.9081 | 14.725 | 12.1578 | 36.454
38 | CREATINE KINASE Gen 2 | Creatine kinase | ±24% | ± 30% | ± 20% | ± 11.3% | ± 30.3% | ± 24% | ±20.4%[a] ±22.6%[f] ± 33.8%[g] | . | Creatine Kinase | Serum/plasma | 15 | 37.9 | June 6th 2022 | 2022-06-08 | 11.25 | 15.2851 | 33.8476 | 7.5 | 10.1901 | 22.5651 | 3.75 | 5.095 | 11.2825 | Creatine kinase (CK) | S- | 28.3% | 30.3% |  |  | 12% ≥ 125 U/L |  | 24 | A: Westgard's recommended choice | Spanish Min
39 | CREATININE | Creatinine | ± 0.3 mg/dL or ± 15% | ± 0.3 mg/dL or ± 15% | ± 0.2 mg/dL or ± 10% | ±12.6% | ± 8.9% | ± 20% | ±6.4%[c] ±7.4%[f] ±11.1%[g] | . | Creatinine | Serum/plasma | 4.4 | 14.1 | June 6th 2022 | 2022-06-08 | 3.3 | 5.539 | 10.984 | 2.2 | 3.6926 | 7.3226 | 1.1 | 1.8463 | 3.6613 | Creatinine | S- | 7.7% | 8.87% |  | 8 umol/L ≤ 100 umol/L | 8% ≥ 100 umol/L |  | 15 | A: Westgard's recommended choice | CLIA
40 | [9 empty] | . | Cystatin C | Serum/plasma | 3.9 | 12.1 | June 6th 2022 | 2022-06-08 | 2.925 | 4.7674 | 9.5936 | 1.95 | 3.1782 | 6.3957 | 0.975 | 1.5891 | 3.1979
41 | DEHYDROEPIANDROSTERONE SULPHAT | [8 empty] | . | Dehydroepiandrosterone sulphate (DHEAS) | Serum/plasma | 5.9 | 21 | June 6th 2022 | 2022-06-08 | 4.425 | 8.1799 | 15.4812 | 2.95 | 5.4533 | 10.3208 | 1.475 | 2.7266 | 5.1604 | Dehydroepiandrosterone sulfate (DHEAS) | S- |  | 13.08% | 19.6% | 1.2 umol/L ≤ 10umol/L | 12% ≥10umol/L |  | 15.5 | C: Algorithm - choice 2 | EFLM_M
42 | [9 empty] | . | Dihydroxytestosterone (DHT) | Serum/plasma | 11.5 | 37.1 | June 6th 2022 | 2022-06-08 | 8.625 | 14.5656 | 28.7968 | 5.75 | 9.7104 | 19.1979 | 2.875 | 4.8552 | 9.5989
43 | [9 empty] | . | Eosinophils | Whole Blood | 14.9 | 65.3 | June 6th 2022 | 2022-06-08 | 11.175 | 25.1169 | 43.5556 | 7.45 | 16.7446 | 29.0371 | 3.725 | 8.3723 | 14.5185
44 | [9 empty] | . | Erythrocytes | Whole Blood | 2.6 | 6.4 | June 6th 2022 | 2022-06-08 | 1.95 | 2.5905 | 5.808 | 1.3 | 1.727 | 3.872 | 0.65 | 0.8635 | 1.936
45 | OESTRADIOL III | Estradiol | ± 26.9% |  | ± 30% | ±13.7% | ± 26.86% | ± 26% | ± 17.3%[f] ± 26.0%[i] | . | Estradiol | Serum/plasma | 15 | 13 | April 20th 2020 | 2022-06-08 | 11.25 | 7.4435 | 26.006 | 7.5 | 4.9624 | 17.3374 | 3.75 | 2.4812 | 8.6687 | Estradiol | S- | 13.4% | 26.86% |  | 25 pmol/L ≤ 100 pmol/L | 25% ≥ 100.0 pmol/L |  | 26.9 | A: Westgard's recommended choice | RICOS
46 | FERRITIN | Ferritin | ± 16.9% |  | ± 20% | ±13.1% | ± 16.9% | ± 21% | ± 13.76% | . | Ferritin | Serum/plasma | 12.8 |  | July 2nd 2020 | 2022-06-08 | 9.6 | 4.8 | 20.64 | 6.4 | 3.2 | 13.76 | 3.2 | 1.6 | 6.88 | Ferritin | S- | 8.4% | 16.9% |  | 4.0ug/L ≤ 27.0ug/L | 15% ≥ 27ug/L |  | 16.9 | A: Westgard's recommended choice | RICOS
47 | [9 empty] | . | Fibroblast growth factor-23 | Serum/plasma | 13.9 | 21.4 | June 6th 2022 | 2022-06-08 | 10.425 | 9.5693 | 26.7705 | 6.95 | 6.3795 | 17.847 | 3.475 | 3.1898 | 8.9235
48 | FOLLICLE STIMULATING HORMONE | Follicle Stimulating Hormone (FSH) | ± 21.2% |  | ± 2 IU/L or 18% | ±11.4% | ± 21.2% | ± 14% | ± 21.2%[f] ± 31.8%[i] | . | Follicle stimulating hormone (FSH) | Serum/plasma | 12.3 | 42.1 | June 6th 2022 | 2022-06-08 | 9.225 | 16.4475 | 31.6688 | 6.15 | 10.965 | 21.1125 | 3.075 | 5.4825 | 10.5563 | Follicle stimulating hormone (FSH) | S- | 10.6% | 21.19% |  | 1.0IU/L ≤ 10.0 IU/L | 10% ≥ 10IU/L |  | 21.2 | A: Westgard's recommended choice | RICOS
49 | [5 empty] | ±22.8% |  |  |  | . | Free kappa light chain | Serum/plasma | 4.8 | 15.3 | May 26th 2020 | 2022-06-08 | 3.6 | 6.0132 | 11.9532 | 2.4 | 4.0088 | 7.9688 | 1.2 | 2.0044 | 3.9844
50 | [5 empty] | ±22.6% |  |  |  | . | Free lambda light chain | Serum/plasma | 4.8 | 17.3 | November 20th 2019 | 2022-06-08 | 3.6 | 6.7326 | 12.6726 | 2.4 | 4.4884 | 8.4484 | 1.2 | 2.2442 | 4.2242
51 | [9 empty] | . | Fructosamine | Serum/plasma | 2.3 | 6.3 | November 20th 2019 | 2022-06-08 | 1.725 | 2.515 | 5.3613 | 1.15 | 1.6767 | 3.5742 | 0.575 | 0.83834 | 1.7871
52 | GAMMA-GLUTAMYL TRANSFERASE | GGT | ± 22.1% |  | ± 5 U/L or ± 15% | ±17.9% | ± 22.1% | ± 22% | ±16.2%[a] ±18.9%[f] ±28.3%[g] | . | g-glutamyl transferase activity | Serum/plasma | 9.1 | 44.4 | June 6th 2022 | 2022-06-08 | 6.825 | 16.9961 | 28.2574 | 4.55 | 11.3307 | 18.8382 | 2.275 | 5.6654 | 9.4191 | γ-glutamyltransferase (GGT) | S- | 11.1% | 22.11% |  | 5 U/L ≤ 40 U/L | 12% ≥ 40 U/L |  | 22.1 | A: Westgard's recommended choice | RICOS
53 | GLUCOSE | Glucose | ± 6 mg/dL or ± 10% | ± 6 mg/dL or ± 10% | ± 8.0% | ±8.5% | ± 7% | ± 12% | ± 6.3%[e] ± 6.5%[f] ± 9.8%[g] | . | Glucose | Serum/plasma | 4.9 | 8.1 | June 6th 2022 | 2022-06-08 | 3.675 | 3.55 | 9.6138 | 2.45 | 2.3667 | 6.4092 | 1.225 | 1.1833 | 3.2046 | Glucose | P- |  | 5.5% |  | 0.4mmol/L ≤ 5.0mmol/L | 8% ≥ 5.0mmol/L |  | 10 | A: Westgard's recommended choice | CLIA
54 | [9 empty] | . | Glycated albumin | Serum/plasma | 1.4 | 5.7 | June 6th 2022 | 2022-06-08 | 1.05 | 2.201 | 3.9335 | 0.7 | 1.4674 | 2.6224 | 0.35 | 0.73368 | 1.3112
55 | [9 empty] | . | Haematocrit (Hct) | Whole Blood | 2.8 | 5.4 | June 6th 2022 | 2022-06-08 | 2.1 | 2.281 | 5.746 | 1.4 | 1.5207 | 3.8307 | 0.7 | 0.76035 | 1.9153
56 | [9 empty] | . | Haemoglobin A1c (IFCC) | Whole Blood | 1.6 | 7.1 | June 6th 2022 | 2022-06-08 | 1.2 | 2.7293 | 4.7093 | 0.8 | 1.8195 | 3.1395 | 0.4 | 0.90976 | 1.5698
57 | RHD3 | HbA1c (NGSP) | ± 6.0% [NGSP] |  | ± 10.0% |  |  |  | ±2.8%[e] ±2.2%[f] ±3.3%[g] | . | Haemoglobin A1c (NGSP) | Whole Blood | 1.1 | 4.8 | June 6th 2022 | 2022-06-08 | 0.825 | 1.8467 | 3.2079 | 0.55 | 1.2311 | 2.1386 | 0.275 | 0.61555 | 1.0693 | Hemoglobin A1 C (NSP) (%) | B- |  | 3% | 4.5% | 0.5 ≤ 10% | 5% ≥ 10% |  | 6 | A: Westgard's recommended choice / B | NGSP
133 | GLYCATED HAEMOGLOBIN (NGSP) | [8 empty] | . | Haemoglobin A1c (NGSP) | Whole Blood | 1.1 | 4.8 | June 6th 2022 | 2022-06-08 | 0.825 | 1.8467 | 3.2079 | 0.55 | 1.2311 | 2.1386 | 0.275 | 0.61555 | 1.0693 | Hemoglobin A1 C (NSP) (%) | B- |  | 3% | 4.5% | 0.5 ≤ 10% | 5% ≥ 10% |  | 6 | A: Westgard's recommended choice / B | NGSP
58 | [9 empty] | . | Haemoglobin (Hb) | Whole Blood | 2.7 | 5.9 | June 6th 2022 | 2022-06-08 | 2.025 | 2.4332 | 5.7744 | 1.35 | 1.6221 | 3.8496 | 0.675 | 0.81106 | 1.9248
59 | HAPTOGLOBIN | Haptoglobin | ± 27.3% |  |  | ±7.9% | ± 27.3% |  | ± 17.1%[f] ± 25.6% [g] | . | Haptoglobin | Serum/plasma | 8.6 | 39 | June 6th 2022 | 2022-06-08 | 6.45 | 14.9764 | 25.6189 | 4.3 | 9.9842 | 17.0792 | 2.15 | 4.9921 | 8.5396 | Haptoglobin | S- |  | 27.3% |  |  |  |  | 27.3 | A: Westgard's recommended choice | RICOS
60 | HDL CHOLESTEROL | Cholesterol, HDL | ± 20% | ± 30% | ± 20% | ±21.1% | ± 11.6% | ± 33% | ±11.1%[d] ±11.1%[f] ±16.6%[g] | . | HDL cholesterol | Serum/plasma | 5.7 | 24.4 | June 6th 2022 | 2022-06-08 | 4.275 | 9.3964 | 16.4501 | 2.85 | 6.2642 | 10.9667 | 1.425 | 3.1321 | 5.4834 | HDL cholesterol | S- |  | 11.63% |  | 0.1 mmol/L ≤ 0.8 mmol/L | 12% ≥ 0.8 mmol/L |  | 20 | A: Westgard's recommended choice | CLIA '19
61 | [9 empty] | . | Human epididymis protein 4 (HE4) | Serum/plasma | 7 | 17.9 | April 19th 2022 | 2022-06-08 | 5.25 | 7.2075 | 15.87 | 3.5 | 4.805 | 10.58 | 1.75 | 2.4025 | 5.29
62 | [9 empty] | . | Hydrogen ion (H+) | Whole Blood | 3.5 | 2 | February 11th 2021 | 2022-06-08 | 2.625 | 1.5117 | 5.8429 | 1.75 | 1.0078 | 3.8953 | 0.875 | 0.50389 | 1.9476
63 | IMMUNOGLOBULIN A | IgA | ± 15% |  | ± 15% | ±7.9% | ± 13.5% | ± 21% | ± 9.8%[f] ±14.7%[g] | . | Immunoglobulin A (IgA) | Serum/plasma | 5.6 | 19.5 | June 6th 2022 | 2022-06-08 | 4.2 | 7.6081 | 14.5381 | 2.8 | 5.072 | 9.692 | 1.4 | 2.536 | 4.846 | Immunoglobulin A | S- | 6.8% | 13.5% |  |  |  |  | 15 | A: Westgard's recommended choice | CLIA '19
64 | IMMUNOGLOBULIN G | IgG | ± 20% |  | ± 20% | ±7.9% | ± 8% | ± 16% | ± 7.3%[f] ±10.9%[f] | . | Immunoglobulin G (IgG) | Serum/plasma | 3.4 | 17.1 | June 6th 2022 | 2022-06-08 | 2.55 | 6.538 | 10.7455 | 1.7 | 4.3587 | 7.1637 | 0.85 | 2.1793 | 3.5818 | Immunoglobulin G | S- |  | 8% |  |  |  |  | 20 | A: Westgard's recommended choice | CLIA '19
65 | IMMUNOGLOBULIN M | IgM | ± 28% |  | ± 20% | ±7.9% | ± 16.8% | ± 28% | ± 17.1%[f] ±25.6%[g] | . | Immunoglobulin M (IgM) | Serum/plasma | 5.9 | 48.5 | January 10th 2020 | 2022-06-08 | 4.425 | 18.3216 | 25.6228 | 2.95 | 12.2144 | 17.0819 | 1.475 | 6.1072 | 8.5409 | Immunoglobulin M | S- | 8.4% | 16.8% |  |  |  |  | 28 | A: Westgard's recommended choice | Spanish Min
66 | INSULIN | Insulin | ± 32.9% |  |  | ±15.3% | ± 32.9% |  | ± 31.5%[f] ± 47.2%[i] | . | Insulin | Serum/plasma | 25.4 | 33.4 | June 6th 2022 | 2022-06-08 | 19.05 | 15.7354 | 47.1679 | 12.7 | 10.4902 | 31.4452 | 6.35 | 5.2451 | 15.7226 | Insulin | S- |  | 32.9% |  | 0.6 mU/L ≤ 5.0mU/L | 12% ≥ 5.0 mU/L |  | 32.9 | A: Westgard's recommended choice | RICOS
67 | Insulin-like growth factor I | [8 empty] | . | Insulin-like growth factor-1 (IGF-1) | Serum/plasma | 9.4 | 27 | October 29th 2019 | 2022-06-08 | 7.05 | 10.7211 | 22.3536 | 4.7 | 7.1474 | 14.9024 | 2.35 | 3.5737 | 7.4512 | Insulin-like growth factor (IGF-1) | S- |  | 24% |  | 3 up to 25 nmol/L | 12% >25nmol/L |  | 24 | C: Algorithm - choice 3 (instead of EFLM_M) | BV_D
68 | [9 empty] | . | Interleukin-6 (IL6) | Serum/plasma | 27.8 | 33.9 | June 6th 2022 | 2022-06-08 | 20.85 | 16.4404 | 50.8429 | 13.9 | 10.9603 | 33.8953 | 6.95 | 5.4801 | 16.9476
69 | IRON | Iron, total | ± 15% | ± 20% | ± 15.0% | ±10% | ± 30.7% | ± 24% | ± 26.7%[f] ± 40.0%[g] | . | Iron | Serum/plasma | 20.6 | 32.3 | June 6th 2022 | 2022-06-08 | 15.45 | 14.3662 | 39.8587 | 10.3 | 9.5775 | 26.5725 | 5.15 | 4.7887 | 13.2862 | Iron | S- | 15.3% | 30.7% |  | 3 umol/L ≤ 25 umol/L | 12% ≥ 25 umol/L |  | 15 | A: Westgard's recommended choice | CLIA '19
70 | LACTATE DEHYDROGENASE | Lactate dehydrogenase (LDH) | ± 15.0% | ± 20% | ± 15.0% | ±12.8% | ± 11.4% | ± 26% | ±7.7%[f] ±11.5%[g] | . | Lactate dehydrogenase (LDH) activity | Serum/plasma | 5.2 | 12.6 | November 18th 2020 | 2022-06-08 | 3.9 | 5.1116 | 11.5466 | 2.6 | 3.4077 | 7.6977 | 1.3 | 1.7039 | 3.8489 | Lactate dehydrogenase (LDH) (L to P) | S- | 5.7% | 11.4% |  | 20 U/L ≤ 250 U/L | 8% ≥ 250 U/L |  | 15 | A: Westgard's recommended choice | CLIA '19
71 | [5 empty] | ±20.6% | ± 11.9% |  |  | . | LDL Cholesterol | Serum/plasma | 8.3 | 26.1 | June 6th 2022 | 2022-06-08 | 6.225 | 10.2705 | 20.5417 | 4.15 | 6.847 | 13.6945 | 2.075 | 3.4235 | 6.8472
72 | [9 empty] | . | Leukocytes | Whole Blood | 10.8 | 16.4 | June 6th 2022 | 2022-06-08 | 8.1 | 7.3638 | 20.7288 | 5.4 | 4.9092 | 13.8192 | 2.7 | 2.4546 | 6.9096
73 | LIPASE | Lipase | ± 37.9% |  |  | ±25.7% | ± 37.9% |  | ±12.6%[a] ±14.2%[f] ±21.3%[g] | . | Lipase | Serum/plasma | 9.2 | 24.8 | June 6th 2022 | 2022-06-08 | 6.9 | 9.9193 | 21.3043 | 4.6 | 6.6129 | 14.2029 | 2.3 | 3.3064 | 7.1014 | Lipase | S- |  | 37.88% |  | 12 U/L ≤ 60 U/L | 20% ≥ 60 U/L |  | 37.9 | A: Westgard's recommended choice | RICOS
74 | LIPOPROTEIN A2 | Lipoprotein (a) | [6 empty] | ±9.35%[f] | . | Lipoprotein (a) | Serum/plasma | 8.8 |  | January 13th 2022 | 2022-06-08 | 6.6 | 3.3 | 14.19 | 4.4 | 2.2 | 9.46 | 2.2 | 1.1 | 4.73 | Lipoprotein (a) | S- |  | 24.1% |  |  |  |  | 24.1 | C: Algorithm - choice 3 (instead of EFLM_M) | BV_D
75 | LUTEINISING HORMONE | Luteinizing Hormone (LH) | ± 27.9% |  | ± 20% | ±12.1% | ± 27.92% |  | ± 28.4%[f] ± 42.6%[i] | . | Luteinising hormone (LH) | Serum/plasma | 22.7 | 30.8 | June 6th 2022 | 2022-06-08 | 17.025 | 14.348 | 42.4393 | 11.35 | 9.5653 | 28.2928 | 5.675 | 4.7827 | 14.1464 | Luteinizing hormone (LH) | S- | 14% | 27.92% |  | 1.5 IU/L ≤ 10.0 IU/L | 15% ≥ 10.0 IU/L |  | 27.9 | A: Westgard's recommended choice | RICOS
76 | [9 empty] | . | Lymphocytes | Whole Blood | 10.8 | 22.6 | June 6th 2022 | 2022-06-08 | 8.1 | 9.393 | 22.758 | 5.4 | 6.262 | 15.172 | 2.7 | 3.131 | 7.586
77 | MAGNESIUM | Magnesium | ± 25% | ± 25% | ± 15.0% | ±11% | ± 4.8% |  | ±4%[b] ±4.0[f] ±6.0%[g] | . | Magnesium | Serum/plasma | 2.8 | 5.7 | June 6th 2022 | 2022-06-08 | 2.1 | 2.3815 | 5.8465 | 1.4 | 1.5876 | 3.8976 | 0.7 | 0.79382 | 1.9488 | Magnesium | S- |  | 4.8% | 7.2% | 0.1 mmol/L ≤ 1.25 mmol/L | 8% ≥ 1.25 mmol/L |  | 11 | C: Algorithm - choice 1 | SOTA
78 | [9 empty] | . | Mean corpuscular haemoglobin concentration (MCHC) | Whole Blood | 0.9 | 1.4 | June 6th 2022 | 2022-06-08 | 0.675 | 0.62412 | 1.7379 | 0.45 | 0.41608 | 1.1586 | 0.225 | 0.20804 | 0.57929
79 | [9 empty] | . | Mean corpuscular haemoglobin (MCH) | Whole Blood | 0.8 | 4.3 | June 6th 2022 | 2022-06-08 | 0.6 | 1.6402 | 2.6302 | 0.4 | 1.0934 | 1.7534 | 0.2 | 0.54672 | 0.87672
80 | [9 empty] | . | Mean corpuscular volume (MCV) | Whole Blood | 0.8 | 3.7 | June 6th 2022 | 2022-06-08 | 0.6 | 1.4196 | 2.4096 | 0.4 | 0.94637 | 1.6064 | 0.2 | 0.47319 | 0.80319
81 | [9 empty] | . | Mean platelet volume (MPV) | Whole Blood | 2.2 | 7 | June 6th 2022 | 2022-06-08 | 1.65 | 2.7516 | 5.4741 | 1.1 | 1.8344 | 3.6494 | 0.55 | 0.9172 | 1.8247
82 | [9 empty] | . | Monocytes | Whole Blood | 13.3 | 22.2 | June 6th 2022 | 2022-06-08 | 9.975 | 9.7047 | 26.1634 | 6.65 | 6.4698 | 17.4423 | 3.325 | 3.2349 | 8.7211
83 | [9 empty] | . | Neuron specific enolase | Serum/plasma | 10.8 | 16.6 | June 6th 2022 | 2022-06-08 | 8.1 | 7.4265 | 20.7915 | 5.4 | 4.951 | 13.861 | 2.7 | 2.4755 | 6.9305
84 | [9 empty] | . | Neutrophils | Whole Blood | 14 | 23.5 | June 6th 2022 | 2022-06-08 | 10.5 | 10.2578 | 27.5828 | 7 | 6.8385 | 18.3885 | 3.5 | 3.4193 | 9.1943
85 | [9 empty] | . | Orosmucoid | Serum/plasma | 7.3 | 24 | June 6th 2022 | 2022-06-08 | 5.475 | 9.4071 | 18.4409 | 3.65 | 6.2714 | 12.2939 | 1.825 | 3.1357 | 6.147
86 | [9 empty] | . | Osteocalcin | Serum/plasma | 8.9 | 32.3 | June 6th 2022 | 2022-06-08 | 6.675 | 12.5639 | 23.5776 | 4.45 | 8.3759 | 15.7184 | 2.225 | 4.188 | 7.8592
87 | [9 empty] | . | Pancreatic Amylase | Serum/plasma | 6.6 | 25.7 | June 6th 2022 | 2022-06-08 | 4.95 | 9.9502 | 18.1177 | 3.3 | 6.6335 | 12.0785 | 1.65 | 3.3167 | 6.0392
88 | PARATHYROID HORMONE | Parathyroid Hormone (PTH I-84) | ± 30.2% |  | ± 30% | ±21.2% | ± 30.2% |  | ± 20.0%[f] ± 30.0%[i] | . | Parathyroid hormone | Serum/plasma | 15.7 | 23.5 | June 6th 2022 | 2022-06-08 | 11.775 | 10.5982 | 30.027 | 7.85 | 7.0655 | 20.018 | 3.925 | 3.5327 | 10.009 | Parathyroid hormone (PTH) | P- |  | 33.43% |  |  |  |  | 30.2 | A: Westgard's recommended choice | RICOS
89 | [9 empty] | . | Partial pressure carbon dioxide | Whole Blood | 4.8 | 5.3 | February 11th 2021 | 2022-06-08 | 3.6 | 2.6814 | 8.6214 | 2.4 | 1.7876 | 5.7476 | 1.2 | 0.89382 | 2.8738
90 | PHOSPHATE | Phosphorous | ± 10.1% |  | ± 0.3 mg/dL or 10% |  | ±10.1 | ± 17% | ± 9.7%[f] ± 14.6%[g] | . | Phosphate | Serum/plasma | 7.7 | 10.7 | June 6th 2022 | 2022-06-08 | 5.775 | 4.9435 | 14.4722 | 3.85 | 3.2956 | 9.6481 | 1.925 | 1.6478 | 4.8241 | Phosphate | S- | 5.1% | 10.11% |  | 0.06 mmol/L ≤ 0.75 mmol/L | 8% ≥ 0.75 mmol/L |  | 10.1 | A: Westgard's recommended choice | RICOS
91 | [9 empty] | . | Plateletcrit | Whole Blood | 6.4 | 13.5 | June 6th 2022 | 2022-06-08 | 4.8 | 5.6026 | 13.5226 | 3.2 | 3.7351 | 9.0151 | 1.6 | 1.8675 | 4.5075
92 | [9 empty] | . | Platelet distribution wide (PDW) | Whole Blood | 3.8 | 12.3 | June 6th 2022 | 2022-06-08 | 2.85 | 4.8276 | 9.5301 | 1.9 | 3.2184 | 6.3534 | 0.95 | 1.6092 | 3.1767
93 | [9 empty] | . | Platelet larger cell ratio (P-LCR) | Whole Blood | 6.7 | 20.9 | July 3rd 2020 | 2022-06-08 | 5.025 | 8.2304 | 16.5216 | 3.35 | 5.4869 | 11.0144 | 1.675 | 2.7435 | 5.5072
94 | POTASSIUM | Potassium | ± 0.3 mmol/L | ± 0.5 mmol/L | ± 0.3 mmol/L | ±5.6% | ± 5.6% | ± 8% | ±4.6%[b] ±4.8%[f] ±7.3%[g] | . | Potassium | Serum/plasma | 4 | 4.1 | June 6th 2022 | 2022-06-08 | 3 | 2.148 | 7.098 | 2 | 1.432 | 4.732 | 1 | 0.716 | 2.366 | Potassium | S- |  | 5.61% | 8.4% | 0.2 mmol/L ≤ 4.0 mmol/L | 5% ≥ 4.0 mmol/L |  | 7.1 | C: Algorithm - choice 2 (instead of SOTA) | EFLM_M
95 | PROLACTIN | Prolactin_I | ± 20% | ± 20% |  | ±17.1% | ± 29.4% | ± 22% | ± 37.37% | . | Prolactin | Serum/plasma | 29.5 | 43 | June 6th 2022 | 2022-06-08 | 22.125 | 19.5549 | 56.0612 | 14.75 | 13.0366 | 37.3741 | 7.375 | 6.5183 | 18.6871 | Prolactin | S- |  | 29.4% |  | 40 mIU/L ≤ 400 mIU/L | 10% ≥ 400 mIU/L |  | 20 | A: Westgard's recommended choice | CLIA
96 | PROSTATE-SPECIFIC AG | Prostate Specific Antigen (PSA) including free (f), conjugated (c), and total (t) | ± 33.6% |  | ± 0.2 ng/dL or 20% | ±11.6% | ± 33.6% | ± 17% | c±21.9%[g] f±17.5%[g] t±16.2%[g] t±16.2%[f] t±24.4%[i] c±21.9%[f] c±32.8%[i] f±17.5%[f] f±26.3%[i] | . | Prostate specific antigen (PSA) | Serum/plasma | 6.8 | 42 | October 20th 2019 | 2022-06-08 | 5.1 | 15.9551 | 24.3701 | 3.4 | 10.6367 | 16.2467 | 1.7 | 5.3184 | 8.1234 | Prostatic specific antigen (PSA) Total | S- | 16.8% | 33.6% |  | 0.4 ug/L ≤ 5.0 ug/L | 8% ≥ 5.0 ug/L |  | 33.6 | A: Westgard's recommended choice | RICOS
97 | [9 empty] | . | Prostate specific antigen (PSA) - Conjugated | Serum/plasma | 8.8 | 57.7 | November 9th 2019 | 2022-06-08 | 6.6 | 21.8877 | 32.7777 | 4.4 | 14.5918 | 21.8518 | 2.2 | 7.2959 | 10.9259
98 | FREE PSA | [8 empty] | . | Prostate specific antigen (PSA) - free | Serum/plasma | 7.1 | 46.2 | October 20th 2019 | 2022-06-08 | 5.325 | 17.5284 | 26.3146 | 3.55 | 11.6856 | 17.5431 | 1.775 | 5.8428 | 8.7715 | Prostatic specific antigen (PSA) Free | S- |  |  |  | 0.2 ug/L ≤ 1.4 ug/L | 15% ≥ 1.4 ug/L |  | 33.6 | A: Westgard's recommended choice | RICOS
99 | TOTAL PROTEIN | Total protein | ± 10% | ± 10% | ± 8.0% | ±8.7% | ± 3.6% | ± 12% | ± 3.4%[b] ± 3.5%[f] ± 5.2%[g] | . | Protein (total) | Serum/plasma | 2.6 | 4.6 | June 6th 2022 | 2022-06-08 | 1.95 | 1.9815 | 5.199 | 1.3 | 1.321 | 3.466 | 0.65 | 0.66049 | 1.733 | Protein | S- |  | 3.63% | 5.4% | 3 g/L ≤ 60 g/L | 5% ≥ 60 g/L |  | 10 | A: Westgard's recommended choice | CLIA
100 | [9 empty] | . | Red cell distribution wide standard deviation (RDW-SD) | Whole Blood | 1.6 | 4.2 | June 6th 2022 | 2022-06-08 | 1.2 | 1.6854 | 3.6654 | 0.8 | 1.1236 | 2.4436 | 0.4 | 0.56181 | 1.2218
101 | Renin | [8 empty] | . | Renin | Serum/plasma | 30.1 | 41.6 | November 18th 2019 | 2022-06-08 | 22.575 | 19.2553 | 56.5041 | 15.05 | 12.8369 | 37.6694 | 7.525 | 6.4184 | 18.8347 | Renin concentration | P- | [6 empty] | 56.5 | C: Algorithm - choice 2 | EFLM_M
102 | [9 empty] | . | Reticulocyte haemoglobin equivalent (RET-He) | Whole Blood | 1.7 | 3.3 | June 6th 2022 | 2022-06-08 | 1.275 | 1.3921 | 3.4958 | 0.85 | 0.92804 | 2.3305 | 0.425 | 0.46402 | 1.1653
103 | [9 empty] | . | Reticulocytes | Whole Blood | 9.7 | 27.1 | June 6th 2022 | 2022-06-08 | 7.275 | 10.7939 | 22.7976 | 4.85 | 7.1959 | 15.1984 | 2.425 | 3.598 | 7.5992
104 | [9 empty] | . | S100 calcium-binding protein B (S100B) | Serum/plasma | 10.2 | 32.7 | November 20th 2019 | 2022-06-08 | 7.65 | 12.8452 | 25.4677 | 5.1 | 8.5635 | 16.9785 | 2.55 | 4.2817 | 8.4892
105 | [9 empty] | . | Selenium | Serum/plasma | 7.7 | 11.3 | June 6th 2022 | 2022-06-08 | 5.775 | 5.1278 | 14.6565 | 3.85 | 3.4185 | 9.771 | 1.925 | 1.7093 | 4.8855
106 | SEX HORMONE BINDING GLOBULIN | Sex Hormone Binding Globulin (SHBG) | ± 20.4% |  |  | ±13.7% | ± 20.4% |  | ± 17.0%[f] ± 25.5%[i] | . | Sexual-hormone-binding-globulin (SHBG) | Serum/plasma | 9.6 | 35.5 | June 6th 2022 | 2022-06-08 | 7.2 | 13.7907 | 25.6707 | 4.8 | 9.1938 | 17.1138 | 2.4 | 4.5969 | 8.5569 | Sex hormone binding globulin (SHBG) | S- |  | 20.42% |  | 6.0 nmol/L ≤ 50.0 nmol/L | 12% ≥ 50.0 nmol/L |  | 20.4 | A: Westgard's recommended choice | RICOS
107 | SODIUM | Sodium | ± 4 mmol/L | ± 4 mmol/L | unchanged | ±3.6% | ± 0.7% | ± | ±0.7%[f] ±1.0%[g] | . | Sodium | Serum/plasma | 0.5 | 0.9 | June 6th 2022 | 2022-06-08 | 0.375 | 0.38609 | 1.0048 | 0.25 | 0.25739 | 0.66989 | 0.125 | 0.1287 | 0.33495 | Sodium | S- |  | 0.73% | 1.1% | 3 mmol/L ≤ 150 mmol/L | 2% ≥ 150 mmol/L |  | 3.6 | C: Algorithm - choice 1 | SOTA
108 | [9 empty] | . | Soluble transferrin receptor (sTfr) | Serum/plasma | 6.9 | 19.1 | May 26th 2020 | 2022-06-08 | 5.175 | 7.6155 | 16.1543 | 3.45 | 5.077 | 10.7695 | 1.725 | 2.5385 | 5.3848
109 | [9 empty] | . | Standard bicarbonate | Serum/plasma | 4.2 | 4.4 | February 11th 2021 | 2022-06-08 | 3.15 | 2.281 | 7.4785 | 2.1 | 1.5207 | 4.9857 | 1.05 | 0.76035 | 2.4928
110 | [9 empty] | . | Suppression of tumorigenicity 2 (ST2) | Serum/plasma | 10.5 | 30.4 | November 20th 2019 | 2022-06-08 | 7.875 | 12.0608 | 25.0546 | 5.25 | 8.0406 | 16.7031 | 2.625 | 4.0203 | 8.3515
111 | [9 empty] | . | Tartrate-resistant acid phosphatase (TRAP) | Serum/plasma | 10.8 | 13.3 | November 25th 2019 | 2022-06-08 | 8.1 | 6.4248 | 19.7898 | 5.4 | 4.2832 | 13.1932 | 2.7 | 2.1416 | 6.5966
112 | TESTOSTERONE | Testosterone including free (f) | ± 23% |  | ± 20 ng/dL or 30% | ±15.5% | ± 13.6% | ± 23% | ± 16.5%[f] ± 24.8%[i] f± 27.3%[f] f± 40.0%[i] | . | Testosterone | Serum/plasma | 12.5 | 21.4 | June 6th 2022 | 2022-06-08 | 9.375 | 9.2937 | 24.7625 | 6.25 | 6.1958 | 16.5083 | 3.125 | 3.0979 | 8.2542 | Testosterone | S- |  | 13.61% |  | 0.4 nmol/L ≤ 2.7 nmol/L | 15% ≥ 2.7 nmol/L |  | 23 | A: Westgard's recommended choice | Spanish Min
113 | [9 empty] | . | Testosterone - free | Serum/plasma | 21.9 | 29 | June 6th 2022 | 2022-06-08 | 16.425 | 13.6276 | 40.7288 | 10.95 | 9.085 | 27.1525 | 5.475 | 4.5425 | 13.5763
114 | [9 empty] | . | Thrombocytes | Whole Blood | 7.5 | 18.6 | June 6th 2022 | 2022-06-08 | 5.625 | 7.5207 | 16.8019 | 3.75 | 5.0138 | 11.2013 | 1.875 | 2.5069 | 5.6006
115 | THYROGLOBULIN v2 | Thyroglobuin | [6 empty] | ± 28.2%[f] ± 42.3%[i] | . | Thyroglobulin | Serum/plasma | 10.5 | 77.2 | June 6th 2022 | 2022-06-08 | 7.875 | 29.2165 | 42.2103 | 5.25 | 19.4777 | 28.1402 | 2.625 | 9.7388 | 14.0701 | Thyroglobulin | S- | 11% | 21.9% |  | 0.2 ug/L ≤ 1.7 ug/L | 12% ≥ 1.7 ug/L |  | 42.2 | C: Algorithm - choice 2 | EFLM_M
116 | THYROID STIMULATING HORMONE | Thyroid Stimulating Hormone (TSH) | ± 23.7% |  | ± 2 mIU/L or 20% | ±11.2% | ± 23.7% | ± 15% | ± 24.6%[f,h] ± 36.9%[i] | . | Thyroid stimulating hormone (TSH) | Serum/plasma | 17.6 | 35.9 | June 6th 2022 | 2022-06-08 | 13.2 | 14.9933 | 36.7733 | 8.8 | 9.9955 | 24.5155 | 4.4 | 4.9978 | 12.2578 | Thyroid stimulating hormone (TSH) | S- | 11.9% | 23.7% |  | 0.1 mU/L ≤ 0.5mU/L | 20% ≥ 0.5 mU/L |  | 23.7 | A: Westgard's recommended choice | RICOS
117 | Free T4 Gen 3 | Thyroxine, Free (FT4) | ± 11.9% |  |  | ±11.9% |  |  | ± 6.2%[h] ± 6.3%[f] ± 9.5%[i] | . | Thyroxine - free (FT4) | Serum/plasma | 4.8 | 7.7 | June 6th 2022 | 2022-06-08 | 3.6 | 3.4026 | 9.3426 | 2.4 | 2.2684 | 6.2284 | 1.2 | 1.1342 | 3.1142 | Thyroxine, free (FT4) | S- |  | 8% |  | 1.5 pmol/L ≤ 12 pmol/L | 12% ≥ 12 pmol/L |  | 11.9 | A: Westgard's recommended choice | SOTA
118 | [5 empty] | ±14.5% | ± 7% | ± 24% |  | . | Thyroxine - total (T4) | Serum/plasma | 6.4 | 12 | June 6th 2022 | 2022-06-08 | 4.8 | 5.1 | 13.02 | 3.2 | 3.4 | 8.68 | 1.6 | 1.7 | 4.34
119 | [5 empty] | ±16.4% | ± 9.22% |  |  | . | Total tri-iodothyronine (T3) | Serum/plasma | 9.4 | 12.2 | June 6th 2022 | 2022-06-08 | 7.05 | 5.7755 | 17.408 | 4.7 | 3.8503 | 11.6053 | 2.35 | 1.9252 | 5.8027
120 | TRANSFERRIN | Transferrin | ± 20% |  |  | ±7.9% | ± 3.8% |  | ±6.8%[f] ±10.2%[f] | . | Transferrin | Serum/plasma | 3.8 | 13.9 | June 6th 2022 | 2022-06-08 | 2.85 | 5.4038 | 10.1063 | 1.9 | 3.6025 | 6.7375 | 0.95 | 1.8013 | 3.3688 | Transferrin | S- |  | 3.8% | 5.7% | 0.2 g/L ≤ 2.5 g/L | 8% >2.50 g/L |  | 20 | A: Westgard's recommended choice | Not stated
121 | [9 empty] | . | Transthyretin | Serum/plasma | 10.9 | 19.1 | January 10th 2020 | 2022-06-08 | 8.175 | 8.2468 | 21.7355 | 5.45 | 5.4978 | 14.4903 | 2.725 | 2.7489 | 7.2452
122 | TRIGLYCERIDE | Triglycerides | ± 25% | ± 25% | ± 15% | ±15.4% | ± 26% | ± 18% | ±28%[d] ±27.0%[f] ±40.6%[g] | . | Triglycerides | Serum/plasma | 20 | 37 | June 6th 2022 | 2022-06-08 | 15 | 15.7723 | 40.5223 | 10 | 10.5149 | 27.0149 | 5 | 5.2574 | 13.5074 | Triglyceride | S- | 13% | 25.99% |  | 0.2 mmol/L ≤ 1.6 mmol/L | 12% ≥ 1.6 mmol/L |  | 25 | A: Westgard's recommended choice | CLIA
123 | TRI-IODO THYRONINE | Triiodothyronine, Free (FT3) | ± 11.6% |  |  | ±13.3% | ± 11.3% |  | ± 6.5%[f,h] ± 9.8%[i] | . | Triiodothyronine - free (FT3) | Serum/plasma | 4.9 | 8.3 | June 6th 2022 | 2022-06-08 | 3.675 | 3.6144 | 9.6782 | 2.45 | 2.4096 | 6.4521 | 1.225 | 1.2048 | 3.2261 | Triiodothyronine, free (FT3) | S- |  | 11.3% |  | 0.7 pmol/L ≤ 3.5 pmol/L | 20% ≥ 3.5 pmol/L |  | 11.6 | A: Westgard's recommended choice | Not stated
124 | [5 empty] | ±19.1% | ± 27.9% |  |  | . | Troponin I, cardiac - high-sensitive (biweekly-monthly sampling) | Serum/plasma | 12 | 35.8 | June 6th 2022 | 2022-06-08 | 9 | 14.1591 | 29.0091 | 6 | 9.4394 | 19.3394 | 3 | 4.7197 | 9.6697
125 | HIGH-SENSITIVITY TROPONIN T | Troponin T | [6 empty] | ± 17.6%[f] ± 26.5%[i] | . | Troponin T, cardiac - high-sensitive (biweekly-monthly sampling) | Serum/plasma | 11.3 | 30.8 | June 6th 2022 | 2022-06-08 | 8.475 | 12.3028 | 26.2866 | 5.65 | 8.2019 | 17.5244 | 2.825 | 4.1009 | 8.7622 | Troponin T | S- |  | 48.9% |  | 0.01 ug/L ≤ 0.05 ug/L | 20% ≥ 0.05 ug/L |  | 10 | B: Published target | Clin Chem
126 | [9 empty] | . | Tumor necrosis factor (TNF) | Serum/plasma | 16.1 | 26.7 | June 6th 2022 | 2022-06-08 | 12.075 | 11.6919 | 31.6157 | 8.05 | 7.7946 | 21.0771 | 4.025 | 3.8973 | 10.5386
127 | [9 empty] | . | Tyrosine | Serum/plasma | 7.5 | 11.9 | November 20th 2019 | 2022-06-08 | 5.625 | 5.2749 | 14.5561 | 3.75 | 3.5166 | 9.7041 | 1.875 | 1.7583 | 4.852
128 | URIC ACID | Uric acid | ± 17% | ± 17% | ± 10% | ±10.6% | ± 12% | ± 17% | ±10.6%[b] ±12.8%[f] ±19.2%[g] | . | Urate | Serum/plasma | 8.2 | 22.4 | June 6th 2022 | 2022-06-08 | 6.15 | 8.9451 | 19.0926 | 4.1 | 5.9634 | 12.7284 | 2.05 | 2.9817 | 6.3642 | Urate | S- | Fso | 11.97% |  | 0.03 mmol/L ≤ 0.38 mmol/L | 8% ≥ 0.38 mmol/L |  | 17 | A: Westgard's recommended choice | CLIA
129 | UREA | Urea Nitrogen | ± 2 mg/dL or ± 9% | ± 2 mg/dL or ± 9% | unchanged | ±12.5% | ± 15.5% | ± 19% | ±16.1%[b] ±17.8%[f] ±26.6%[g] | . | Urea | Serum/plasma | 13.9 | 20.9 | June 6th 2022 | 2022-06-08 | 10.425 | 9.4126 | 26.6138 | 6.95 | 6.275 | 17.7425 | 3.475 | 3.1375 | 8.8713 | Urea | S- | 7.8% | 15.55% |  | 0.5 mmol/L ≤ 4.0 mmol/L | 12% ≥ 4.0 mmol/L |  | 9 | A: Westgard's recommended choice | CLIA
130 | [9 empty] | . | Vitamin A | Serum/plasma | 6.1 | 21 | June 6th 2022 | 2022-06-08 | 4.575 | 8.2005 | 15.7493 | 3.05 | 5.467 | 10.4995 | 1.525 | 2.7335 | 5.2498
131 | [9 empty] | . | Vitamin E | Serum/plasma | 7 | 17.3 | June 6th 2022 | 2022-06-08 | 5.25 | 6.9984 | 15.6609 | 3.5 | 4.6656 | 10.4406 | 1.75 | 2.3328 | 5.2203
132 | [9 empty] | . | Zinc | Serum/plasma | 8.5 | 8.8 | June 6th 2022 | 2022-06-08 | 6.375 | 4.588 | 15.1068 | 4.25 | 3.0587 | 10.0712 | 2.125 | 1.5293 | 5.0356
134 | ALKALINE PHOSPHATASE | Alkaline phosphatase | ±14.5% | ± 30% | ± 20% | ±18.5% | ± 12.0% | ± 31% | ±10.7%[a] ±14.5%[f] (bone) ±21.7%[g] (bone) ±25.8%[g] (liver) | . | [15 empty] | Alkaline phosphatase (ALP) | S- |  | 12.04% |  | 15 U/L ≤ 125 U/L | 12% ≥ 125 U/L |  | 14.5 | A: Westgard's recommended choice | Not stated
135 | BICARBONATE | Bicarbonate | ± 5 mm Hg or ± 8% | ± 5 mm Hg or ± 8% | ± 20% | ±17.3% | ± 4.86% |  | ± 7.3%[g] | . | [15 empty] | CO2, total | B- |  |  |  | 2.0 umol/L ≤ 20 umol/L | 10% ≥ 20 umol/L |  | 17.3 | C: Algorithm - choice 1 | SOTA
144 | AMMONIA | [8 empty] | . | [15 empty] | Ammonia | WB- |  |  |  | 5umol/L ≤50umol/L | 20% ≥50umol/L |  | 20 | C: Algorithm - choice 4 | RCPA_U
136 | CONJUGATED BILIRUBIN | Bilirubin, direct | ± 44.5% |  |  | ±25.8% | ± 44.5% |  | ±33.4%[b] | . | [15 empty] | Bilirubin conjugated | S- | 22.3% | 44.5% |  | 3.0 umol/L ≤ 15 umol/L | 20% ≥ 15 umol/L |  | 44.5 | A: Westgard's recommended choice | RICOS
145 | Osmolality | [8 empty] | . | [15 empty] | Osmolality | S- |  | 1.5% | 2.3% | 8 mmol/kg ≤ 266 mmol/kg | 3% ≥ 266 mmol/kg |  | 3 | C: Algorithm - choice 4 (instead of BV_D) | RCPA_U
137 | C-peptide | C-peptide | ± 20.8% |  |  | ±14.8% | ± 20.8% |  |  | . | [15 empty] | C Peptide | S- | 10.4% | 20.8% |  | 0.15 up to 1.25nmol/L | 12% >1.25 nmol/L |  | 20.8 | A: Westgard's recommended choice | RICOS
138 | SERUM FOLATE III | Folate | ± 39% |  | ± 1 ng/mL or 30% | ±23.9% | ± 39% |  |  | . | [15 empty] | Folate | S- | 19.5% | 39% |  | 1.5nmol/L ≤ 6.0nmol/L | 25% ≥ 6.0nmol/L |  | 39 | A: Westgard's recommended choice | RICOS
139 | BETA-HCG | Human Chorionic Gonadotropin (BhCG) | ± 18% or positive/negative |  | ± 18% or positive/negative |  |  |  |  | . | [15 empty] | β-hCG | S- |  |  |  | 1.0 U/L ≤ 10 U/L | 10% ≥ 10 U/L |  | 18 | A: Westgard's recommended choice | CLIA '19
140 | LACTATE | Lactate | ± 30.4% |  |  | ±9.1% | ± 30.4% |  | ±36.2%[e] | . | [15 empty] | Lactate | B- | 15.2% | 30.4% |  | 0.5 mmol/L ≤ 4.0 mmol/L | 12% ≥ 4.0 mmol/L |  | 30.4 | A: Westgard's recommended choice | RICOS
146 | proBNP | [8 empty] | . | [15 empty] | N-terminal (NT)-proBNP | S- |  | 13% |  | 25ng/L≤125ng/L | 20%≥125ng/L |  | 20 | C: Algorithm - choice 4 (instead of BV_D) | RCPA_U
147 | PROCALCITONIN (SENSITIVE) | [8 empty] | . | [15 empty] | Procalcitonin | S- | [6 empty] | 20.3 | B: Published target | Clin Chem
141 | LITHIUM | Lithium | ± 15% | ± 0.3 mmol/L or 20% | ± 15% | ±11.7% |  | ± 18% |  | . | [15 empty] | Lithium | S- |  |  |  | 0.2 mmol/L |  |  | 15 | A: Westgard's recommended choice | CLIA '19
148 | RHEUMATOID FACTOR | [8 empty] | . | [15 empty] | Rheumatoid factor | S- |  | 13.5% |  |  |  |  | 13.5 | C: Algorithm - choice 3 | BV_D
142 | PROGESTERONE III | Progesterone | ± 25% | ± 25% |  | ±15.6% |  | ± 26% |  | . | [15 empty] | Progesterone | S- |  |  |  | 2.0 nmol/L ≤ 10.0 nmol/L | 15% ≥ 10.0 nmol/L |  | 25 | A: Westgard's recommended choice | CLIA
143 | VITAMIN B12 II | Vitamin B12 | ± 30% | ± 30% | ± 25% | ±17.3% | ± 30% |  |  | . | [15 empty] | Vitamin B12 | S- |  |  |  | 18 pmol/L ≤ 120 pmol/L | 15% ≥ 120 pmol/L |  | 30 | A: Westgard's recommended choice | CLIA
149 | ADRENOCORTICOTROPIC HORMONE | [8 empty] | . | [15 empty] | ACTH |  |  |  |  | +-2.0 up to 20 pmol/L | 10% >20 pmol/L |  | 10 | C: Algorithm - choice 4 | RCPA_L and U
150 | Human growth hormone | [8 empty] | . | [15 empty] | Growth Hormone | S- |  |  |  | +-1 up to ≤ 7 mU/L | 15% ≥ 7 mU/L |  | 15 | C: Algorithm - choice 4 | RCPA_U
151 | PARACETAMOL 2 | [8 empty] | . | [15 empty] | Acetaminophen (Paracetamol) | S- |  |  |  | 20 umol/L ≤ 200 umol/L | 10% ≥ 200 umol/L |  | 10 | C: Algorithm - choice 4 | RCPA_U
152 | SODIUM URINE | [6 empty] | 10 |  | . | [15 empty] | Sodium, output, 24 h. | U- | 16% | 32% |  | 2 mmol/L ≤ 20 mmol/L | 10% ≥ 20 mmol/L |  | 32 | C: Algorithm - choice 3 | BV_D
153 | POTASSIUM URINE | [6 empty] | 12 |  | . | [15 empty] | Potassium, output | U- | 14.2% | 28.4% |  | 2 mmol/L ≤ 20.0 mmol/L | 10% ≥ 20.0 mmol/L |  | 28.4 | C: Algorithm - choice 3 | BV_D
154 | CHLORIDE URINE | [6 empty] | 12 |  | . | [15 empty] | Chloride | U- |  |  |  | 2 mmol/L ≤ 20 mmol/L | 10% ≥ 20 mmol/L |  | 12 | A: Westgard's recommended choice | Spanish Min
155 | UREA URINE | [6 empty] | 19 |  | . | [15 empty] | Urea, output, 24h | U- | 11% | 22.1% |  | 20 mmol/L ≤ 200 mmol/L | 10% ≥ 200 mmol/L |  | 22.1 | C: Algorithm - choice 3 | BV_D
156 | CREATININE URINE | [6 empty] | 15 |  | . | [15 empty] | Creatinine, output, 24h | U- |  | 15.4% |  | 0.5 mmol/L ≤ 5 mmol/L | 10% ≥ 5 mmol/L |  | 15.4 | C: Algorithm - choice 3 | BV_D
157 | Urine osmolality | [8 empty] | . | [15 empty] | Osmolality, first morning | U- | 19.7% | 39.5% |  | 6 mmol/kg ≤ 300 mmol/kg | 2% ≥ 300 mmol/kg |  | 2 | C: Algorithm - choice 4 (instead of BV_D) | RCPA_U
158 | CALCIUM URINE | [6 empty] | 30 |  | . | [15 empty] | Calcium, output, 24h | U- | 15.5% | 31% |  |  |  |  | 31 | C: Algorithm - choice 3 | BV_D
159 | MAGNESIUM URINE | [8 empty] | . | [15 empty] | Magnesium, output, 24h | U- |  | 45% |  | 0.2 mmolL |  |  | 45 | C: Algorithm - choice 3 | BV_D
160 | INORGANIC PHOSPHATE URINE | [6 empty] | 16 |  | . | [15 empty] | Phosphate, output, 24h | U- | 11% | 22.1% |  | 2.5mmol/L |  |  | 22.1 | C: Algorithm - choice 3 | BV_D
161 | URIC ACID URINE | [6 empty] | 15 |  | . | [15 empty] | Urate, output, 24h | U- | 9.7% | 19.39% |  | 0.3mmol/L |  |  | 19.4 | C: Algorithm - choice 3 | BV_D
162 | PROTEIN URINE | [6 empty] | 34 |  | . | [15 empty] | Protein, output, 24h | U- | 20% | 40% |  | 0.1 g/L ≤ 1.0 g/L | 10% ≥ 1.0 g/L |  | 40 | C: Algorithm - choice 3 | BV_D
163 | ALBUMIN URINE | [6 empty] | 38 |  | . | [15 empty] | Albumin, output, night urine | U- |  | 40.6% |  | 4.0 g/L ≤ 20 g/L | 20% ≥ 20 g/L |  | 40.6 | C: Algorithm - choice 3 | BV_D
164 | PROTEIN CSF | [8 empty] | . | [15 empty] | Protein | CSF- |  |  |  | 0.02g/L ≤ 0.45g/L | 5% ≥ 0.45g/L | 23 | 23 | C: Algorithm - choice 5 | Rilibak
165 | ALBUMIN CSF | [8 empty] | . | [15 empty] | Albumin, CSF | CSF- |  |  |  | 0.02g/L≤0.45g/L | 5% ≥0.45g/L | 23 | 23 | C: Algorithm - choice 5 | Rilibak
166 | IGG CSF | [8 empty] | . | [15 empty] | Immunoglobulin G | CSF- |  |  |  | 0.02g/L ≤ 0.1g/L | 20% ≥0.1g/L |  | 20 | C: Algorithm - choice 4 | RCPA_U
167 | ADENOSINE DEAMINASE FLUID | [8 empty] | . | [15 empty] | Adenosine deaminase (ADA) | S- |  | 16.7% |  |  |  |  | 16.7 | C: Algorithm - choice 3 | BV_D
168 | ANTI-THYROGLOBULIN AB | [8 empty] | . | [15 empty] | Thyroglobulin antibody | S- |  | 27.6% |  |  |  |  | 27.6 | C: Algorithm - choice 3 | BV_D
169 | ANTI-THYROID PEROXIDASE AB | [8 empty] | . | [15 empty] | Thyroid peroxidase antibody | S- |  | 46.2% |  |  |  |  | 46.2 | C: Algorithm - choice 3 | BV_D

## Cobas assays
Monthly Total Error vs Total allowable error
Laboratory | GSH | [5 empty] | Department |  | C17 CHEM PATH
Date | [6 empty] | Done by | [9 empty] | Approved by |  |  | Date completed: 3/8/2022
Instrument | Analyte | Control Sample | Lot number | Units | Target Value | # | TEa% | TEa (in units) | * | Mean | SD | TE (in units) | N | CV% | %Bias | SEc | Sigma | TE < TAE - Pass/ > TAE -Fail | Pathologist Comment: | IQC Officer comment: E. Samsodien
| ADENOSINE DEAMINASE FLUID |  |  | U/l |  | 2 | 16.7 |  | C3
| ADENOSINE DEAMINASE FLUID |  |  | U/l |  | 2 | 16.7 |  | C3
| ADENOSINE DEAMINASE FLUID |  |  | U/l |  | 2 | 16.7 |  | C3
| ADRENOCORTICOTROPIC HORMONE |  |  | pmol/l |  | 4 | - | 2 | C4
| ADRENOCORTICOTROPIC HORMONE |  |  | pmol/l |  | 4 | 10 |  | C4
| ALANINE TRANSAMINASE |  |  | U/l |  | 1 | 16.1 |  | A
| ALANINE TRANSAMINASE |  |  | U/l |  | 1 | 16.1 |  | A
| ALANINE TRANSAMINASE |  |  | U/l |  | 1 | 16.1 |  | A
| ALANINE TRANSAMINASE |  |  | U/l |  | 1 | 16.1 |  | A
| ALBUMIN |  |  | g/l |  | 1 | 10 |  | A
| ALBUMIN |  |  | g/l |  | 1 | 10 |  | A
| ALBUMIN |  |  | g/l |  | 1 | 10 |  | A
| ALBUMIN |  |  | g/l |  | 2 | 10 |  | A
| ALBUMIN CSF |  |  | mg/l |  | 2 | 23 |  | C5
| ALBUMIN CSF |  |  | mg/l |  | 1 | 23 |  | C5
| ALBUMIN URINE |  |  | mg/l |  | 2 | 40.6 |  | C3
| ALBUMIN URINE |  |  | mg/l |  | 2 | 40.6 |  | C3
| ALKALINE PHOSPHATASE |  |  | U/l |  | 1 | 14.5 |  | A
| ALKALINE PHOSPHATASE |  |  | U/l |  | 1 | 14.5 |  | A
| ALKALINE PHOSPHATASE |  |  | U/l |  | 1 | 14.5 |  | A
| ALKALINE PHOSPHATASE |  |  | U/l |  | 1 | 14.5 |  | A
| ALPHA 1-ANTITRYPSIN |  |  | g/l |  | 2 | 9.6 |  | A
| ALPHA 1-ANTITRYPSIN |  |  | g/l |  | 2 | 9.6 |  | A
| ALPHA-FETO PROTEIN |  |  | ng/ml |  | 2 | 21.9 |  | A
| ALPHA-FETO PROTEIN |  |  | ng/ml |  | 2 | 21.9 |  | A
| AMMONIA |  |  | umol/l |  | 2 | 20 |  | C4
| AMMONIA |  |  | umol/l |  | 2 | 20 |  | C4
| Anti-Mullerian Hormone |  |  | pmol/l |  | C2 | 20.6
| Anti-Mullerian Hormone |  |  | pmol/l |  | C2 | 20.6
| ANTI-THYROGLOBULIN AB |  |  | IU/ml |  | 1 | 27.6 |  | C3
| ANTI-THYROGLOBULIN AB |  |  | IU/ml |  | 1 | 27.6 |  | C3
| ANTI-THYROID PEROXIDASE AB |  |  | IU/ml |  | 2 | 46.2 |  | C3
| ANTI-THYROID PEROXIDASE AB |  |  | IU/ml |  | 2 | 46.2 |  | C3
| APOPROTEIN A |  |  | g/l |  | 2 | 11.3 |  | A
| APOPROTEIN A |  |  | g/l |  | 2 | 11.3 |  | A
| APOPROTEIN B |  |  | g/l |  | 2 | 11.6 |  | A
| APOPROTEIN B |  |  | g/l |  | 2 | 11.6 |  | A
| ASPARTATE TRANSAMINASE |  |  | U/l |  | 1 | 16.7 |  | A
| ASPARTATE TRANSAMINASE |  |  | U/l |  | 1 | 16.7 |  | A
| ASPARTATE TRANSAMINASE |  |  | U/l |  | 1 | 16.7 |  | A
| ASPARTATE TRANSAMINASE |  |  | U/l |  | 1 | 16.7 |  | A
| BETA-HCG |  |  | U/l |  | 2 | 18 |  | A
| BETA-HCG |  |  | U/l |  | 2 | 18 |  | A
| BETA-HCG |  |  | U/l |  | 2 | 18 |  | A
| BETA-HCG |  |  | U/l |  | 2 | 18 |  | A
| BICARBONATE |  |  | mmol/l |  | 1 | 17.3 | 2 | C1
| BICARBONATE |  |  | mmol/l |  | 2 | 17.3 |  | C1
| CA 125 |  |  | U/ml |  | 2 | 35.4 |  | A
| CA 125 |  |  | U/ml |  | 2 | 35.4 |  | A
| CA 19-9 |  |  | U/ml |  | 1 | 46 |  | A
| CA 19-9 |  |  | U/ml |  | 2 | 46 |  | A
| CALCIUM |  |  | mmol/l |  | 2 | 8.3 |  | C1
| CALCIUM |  |  | mmol/l |  | 2 | 8.3 |  | C1
| CALCIUM |  |  | mmol/l |  | 2 | 8.3 |  | C1
| CALCIUM |  |  | mmol/l |  | 2 | 8.3 |  | C1
| CALCIUM URINE |  |  | mmol/l |  | 1 | 31 |  | C3
| CALCIUM URINE |  |  | mmol/l |  | 1 | 31 |  | C3
| CARCINOEMBRYONIC AG |  |  | mmol/l |  | 2 | 24.7 |  | A
| CARCINOEMBRYONIC AG |  |  | mmol/l |  | 1 | 24.7 |  | A
| CHLORIDE |  |  | mmol/l |  | 3 | 5 |  | A
| CHLORIDE |  |  | mmol/l |  | 2 | 5 |  | A
| CHLORIDE |  |  | mmol/l |  | 3 | 5 |  | A
| CHLORIDE |  |  | mmol/l |  | 2 | 5 |  | A
| CHLORIDE URINE |  |  | mmol/l |  | 1 | 12 |  | A
| CHLORIDE URINE |  |  | mmol/l |  | 1 | 12 |  | A
| COMPLEMENT C3 |  |  | g/l |  | 2 | 15 |  | A
| COMPLEMENT C3 |  |  | g/l |  | 2 | 15 |  | A
| COMPLEMENT C4 |  |  | g/l |  | 3 | 16 |  | A
| COMPLEMENT C4 |  |  | g/l |  | 3 | 16 |  | A
| CONJUGATED BILIRUBIN |  |  | umol/l |  | 1 | 44.5 |  | A
| CONJUGATED BILIRUBIN |  |  | umol/l |  | 3 | 44.5 |  | A
| CONJUGATED BILIRUBIN |  |  | umol/l |  | 3 | 44.5 |  | A
| CONJUGATED BILIRUBIN |  |  | umol/l |  | 1 | 44.5 |  | A
| CONJUGATED BILIRUBIN |  |  | umol/l |  | 1 | 44.5 |  | A
| CONJUGATED BILIRUBIN |  |  | umol/l |  | 1 | 44.5 |  | A
| CORTISOL II |  |  | nmol/l |  | 1 | 22.8 |  | A
| CORTISOL II |  |  | nmol/l |  | 2 | 22.8 |  | A
| C-peptide | [5 empty] | 20.8
| C-peptide | [5 empty] | 20.8
| C-REACTIVE PROTEIN |  |  | mg/l |  | 2 | 50.7 |  | A
| C-REACTIVE PROTEIN |  |  | mg/l |  | 2 | 50.7 |  | A
| C-REACTIVE PROTEIN |  |  | mg/l |  | 2 | 50.7 |  | A
| C-REACTIVE PROTEIN |  |  | mg/l |  | 2 | 50.7 |  | A
| CREATINE KINASE Gen 2 |  |  | U/l |  | 2 | 24 |  | A
| CREATINE KINASE Gen 2 |  |  | U/l |  | 2 | 24 |  | A
| CREATININE |  |  | umol/l |  | 1 | 15 |  | A
| CREATININE |  |  | umol/l |  | 1 | 15 |  | A
| CREATININE |  |  | umol/l |  | 1 | 15 |  | A
| CREATININE |  |  | umol/l |  | 1 | 15 |  | A
| CREATININE URINE |  |  | mmol/l |  | 9 | 15.4 |  | C3
| CREATININE URINE |  |  | mmol/l |  | 9 | 15.4 |  | C3
| DEHYDROEPIANDROSTERONE SULPHAT |  |  | umol/l |  | 2 | 15.5 |  | C2
| DEHYDROEPIANDROSTERONE SULPHAT |  |  | umol/l |  | 2 | 15.5 |  | C2
| FERRITIN |  |  | ng/ml |  | 1 | 16.9 |  | A
| FERRITIN |  |  | ng/ml |  | 1 | 16.9 |  | A
| FOLLICLE STIMULATING HORMONE |  |  | U/l |  | 2 | 21.2 |  | A
| FOLLICLE STIMULATING HORMONE |  |  | U/l |  | 2 | 21.2 |  | A
| FREE PSA |  |  | ng/ml |  | 2 | 33.6 |  | A
| FREE PSA |  |  | ng/ml |  | 2 | 33.6 |  | A
| Free T4 Gen 3 |  |  | pmol/l |  | 1 | 11.9 |  | A
| Free T4 Gen 3 |  |  | pmol/l |  | 2 | 11.9 |  | A
| GAMMA-GLUTAMYL TRANSFERASE |  |  | U/l |  | 2 | 22.1 |  | A
| GAMMA-GLUTAMYL TRANSFERASE |  |  | U/l |  | 2 | 22.1 |  | A
| GAMMA-GLUTAMYL TRANSFERASE |  |  | U/l |  | 2 | 22.1 |  | A
| GAMMA-GLUTAMYL TRANSFERASE |  |  | U/l |  | 2 | 22.1 |  | A
| GLUCOSE |  |  | mmol/l |  | 1 | 10 |  | A
| GLUCOSE |  |  | mmol/l |  | 1 | 10 |  | A
| GLUCOSE |  |  | mmol/l |  | 1 | 10 |  | A
| GLUCOSE |  |  | mmol/l |  | 1 | 10 |  | A
| GLYCATED HAEMOGLOBIN (NGSP) |  |  | % |  | 2 | 6 |  | B
| GLYCATED HAEMOGLOBIN (NGSP) |  |  | % |  | 2 | 6 |  | B
| GLYCATED HAEMOGLOBIN (NGSP) |  |  | % |  | 2 | 6 |  | B
| GLYCATED HAEMOGLOBIN (NGSP) |  |  | % |  | 2 | 6 |  | B
| HAPTOGLOBIN |  |  | g/l |  | 2 | 27.3 |  | A
| HAPTOGLOBIN |  |  | g/l |  | 2 | 27.3 |  | A
| HDL CHOLESTEROL |  |  | mmol/l |  | 2 | 20 |  | A
| HDL CHOLESTEROL |  |  | mmol/l |  | 2 | 20 |  | A
| HIGH-SENSITIVITY TROPONIN T |  |  | ng/l |  | 3 | 17.6 |  | B
| HIGH-SENSITIVITY TROPONIN T |  |  | ng/l |  | 1 | 17.6 |  | B
| HIGH-SENSITIVITY TROPONIN T |  |  | ng/l |  | 1 | 17.6 |  | C
| HIGH-SENSITIVITY TROPONIN T |  |  | ng/l |  | 1 | 17.6 |  | C
| IGG CSF |  |  | mg/l |  | 1 | 20 |  | C4
| IGG CSF |  |  | mg/l |  | 1 | 20 |  | C4
| IMMUNOGLOBULIN A |  |  | g/l |  | 2 | 15 |  | A
| IMMUNOGLOBULIN A |  |  | g/l |  | 2 | 15 |  | A
| IMMUNOGLOBULIN G |  |  | g/l |  | 1 | 20 |  | A
| IMMUNOGLOBULIN G |  |  | g/l |  | 1 | 20 |  | A
| IMMUNOGLOBULIN M |  |  | g/l |  | 2 | 28 |  | A
| IMMUNOGLOBULIN M |  |  | g/l |  | 1 | 28 |  | A
| INORGANIC PHOSPHATE URINE |  |  | mmol/l |  | 1 | 22.1 |  | C3
| INORGANIC PHOSPHATE URINE |  |  | mmol/l |  | 1 | 22.1 |  | C3
| INSULIN |  |  | uU/ml |  | 2 | 32.9 |  | A
| INSULIN |  |  | uU/ml |  | 2 | 32.9 |  | A
| IRON |  |  | umol/l |  | 2 | 15 |  | A
| IRON |  |  | umol/l |  | 2 | 15 |  | A
| LACTATE |  |  | mmol/l |  | 2 | 30.4 |  | A
| LACTATE |  |  | mmol/l |  | 2 | 30.4 |  | A
| LACTATE DEHYDROGENASE |  |  | U/l |  | 2 | 15 |  | A
| LACTATE DEHYDROGENASE |  |  | U/l |  | 2 | 15 |  | A
| LACTATE DEHYDROGENASE |  |  | U/l |  | 2 | 15 |  | A
| LACTATE DEHYDROGENASE |  |  | U/l |  | 2 | 15 |  | A
| LIPASE |  |  | U/l |  | 2 | 37.9 |  | A
| LIPASE |  |  | U/l |  | 2 | 37.9 |  | A
| LIPOPROTEIN A2 |  |  | nmol/l |  | 2 | 24.1 |  | C3
| LIPOPROTEIN A2 |  |  | nmol/l |  | 2 | 24.1 |  | C3
| LITHIUM |  |  | mmol/l |  | 2 | 15 |  | A
| LITHIUM |  |  | mmol/l |  | 2 | 15 |  | A
| LUTEINISING HORMONE |  |  | U/l |  | 2 | 27.9 |  | A
| LUTEINISING HORMONE |  |  | U/l |  | 2 | 27.9 |  | A
| MAGNESIUM |  |  | mmol/l |  | 2 | 11 |  | C1
| MAGNESIUM |  |  | mmol/l |  | 2 | 11 |  | C1
| MAGNESIUM |  |  | mmol/l |  | 2 | 11 |  | C1
| MAGNESIUM |  |  | mmol/l |  | 2 | 11 |  | C1
| MAGNESIUM URINE |  |  | mmol/l |  | 1 | 45 |  | C3
| MAGNESIUM URINE |  |  | mmol/l |  | 1 | 45 |  | C3
| OESTRADIOL III |  |  | pmol/l |  | 1 | 26.9 |  | A
| OESTRADIOL III |  |  | pmol/l |  | 2 | 26.9 |  | A
| Osmolality |  |  | mosmol/kg |  | 2 | 3 |  | C4
| Osmolality |  |  | mosmol/kg |  | 2 | 3 |  | C4
| PARACETAMOL 2 |  |  | umol/l |  | 2 | 10 |  | C4
| PARACETAMOL 2 |  |  | umol/l |  | 4 | 10 |  | C4 | [10 empty] | One outlier result deleted (407.6) - brings performance to within allowable error.
| PARATHYROID HORMONE |  |  | pmol/l |  | 2 | 30.2 |  | A
| PARATHYROID HORMONE |  |  | pmol/l |  | 2 | 30.2 |  | A
| PHOSPHATE |  |  | mmol/l |  | 1 | 10.1 |  | A
| PHOSPHATE |  |  | mmol/l |  | 1 | 10.1 |  | A
| PHOSPHATE |  |  | mmol/l |  | 1 | 10.1 |  | A
| PHOSPHATE |  |  | mmol/l |  | 1 | 10.1 |  | A
| POTASSIUM |  |  | mmol/l |  | 2 | 7.1 |  | C2
| POTASSIUM |  |  | mmol/l |  | 2 | 7.1 |  | C2
| POTASSIUM |  |  | mmol/l |  | 2 | 7.1 |  | C2
| POTASSIUM |  |  | mmol/l |  | 2 | 7.1 |  | C2
| POTASSIUM URINE |  |  | mmol/l |  | 1 | 28.4 |  | C3
| POTASSIUM URINE |  |  | mmol/l |  | 1 | 28.4 |  | C3
| proBNP |  |  | pg/ml |  | 2 | 20 |  | C4
| proBNP |  |  | pg/ml |  | 2 | 20 |  | C4
| PROCALCITONIN (SENSITIVE) |  |  | ng/ml |  | 2 | 20.3 |  | B
| PROCALCITONIN (SENSITIVE) |  |  | ng/ml |  | 2 | 20.3 |  | B
| PROGESTERONE III |  |  | nmol/L |  | 4 | 25 |  | A
| PROGESTERONE III |  |  | ng/ml |  | 2 | 25 |  | A
| PROLACTIN |  |  | g/l |  | 2 | 20 |  | A
| PROLACTIN |  |  | g/l |  | 2 | 20 |  | A
| PROSTATE-SPECIFIC AG |  |  | g/l |  | 2 | 33.6 |  | A
| PROSTATE-SPECIFIC AG |  |  | g/l |  | 1 | 33.6 |  | A
| PROTEIN CSF |  |  | g/l |  | 2 | 23 |  | C5
| PROTEIN CSF |  |  | g/l |  | 2 | 23 |  | C5
| PROTEIN URINE |  |  | mIU/l |  | 2 | 20 |  | C3
| PROTEIN URINE |  |  | mIU/l |  | 2 | 20 |  | C3
| RHD3 |  |  | nmol/l |  | 2 | 6 |  | B
| RHD3 |  |  | nmol/l |  | 2 | 6 |  | B
| RHEUMATOID FACTOR |  |  | U/l |  | 2 | 13.5 |  | C3
| RHEUMATOID FACTOR |  |  | U/l |  | 2 | 13.5 |  | C3
| SERUM FOLATE III |  |  | nmol/l |  | 1 | 39 |  | A
| SERUM FOLATE III |  |  | nmol/l |  | 1 | 39 |  | A
| SEX HORMONE BINDING GLOBULIN |  |  | mmol/l |  | 2 | 20.4 |  | A
| SEX HORMONE BINDING GLOBULIN |  |  | mmol/l |  | 2 | 20.4 |  | A
| SODIUM |  |  | mmol/l |  | 2 | 3.6 |  | C1
| SODIUM |  |  | mmol/l |  | 2 | 3.6 |  | C1
| SODIUM |  |  | mmol/l |  | 2 | 3.6 |  | C1
| SODIUM |  |  | mmol/l |  | 2 | 3.6 |  | C1
| SODIUM URINE |  |  | nmol/l |  | 1 | 32 |  | C3
| SODIUM URINE |  |  | nmol/l |  | 1 | 32 |  | C3
| TESTOSTERONE |  |  | ng/ml |  | 2 | 23 |  | A
| TESTOSTERONE |  |  | ng/ml |  | 2 | 23 |  | A
| THYROGLOBULIN v2 |  |  | uIU/ml |  | 2 | 42.2 |  | C2
| THYROGLOBULIN v2 |  |  | uU/ml |  | 2 | 42.2 |  | C2
| THYROID STIMULATING HORMONE |  |  | uU/ml |  | 2 | 23.7 |  | A
| THYROID STIMULATING HORMONE |  |  | uU/ml |  | 2 | 23.7 |  | A
| THYROID STIMULATING HORMONE |  |  | uU/ml |  | 2 | 23.7 |  | A
| THYROID STIMULATING HORMONE |  |  | uU/ml |  | 2 | 23.7 |  | A
| TOTAL BILIRUBIN |  |  | umol/l |  | 1 | 20 |  | A
| TOTAL BILIRUBIN |  |  | umol/l |  | 1 | 20 |  | A
| TOTAL BILIRUBIN |  |  | umol/l |  | 1 | 20 |  | A
| TOTAL BILIRUBIN |  |  | umol/l |  | 1 | 20 |  | A
| TOTAL BILIRUBIN |  |  | umol/l |  | 1 | 20 |  | A
| TOTAL BILIRUBIN |  |  | umol/l |  | 1 | 20 |  | A
| TOTAL CHOLESTEROL |  |  | mmol/l |  | 2 | 10 |  | A
| TOTAL CHOLESTEROL |  |  | mmol/l |  | 2 | 10 |  | A
| TOTAL CHOLESTEROL |  |  | mmol/l |  | 2 | 10 |  | A
| TOTAL CHOLESTEROL |  |  | mmol/l |  | 2 | 10 |  | A
| TOTAL PROTEIN |  |  | g/l |  | 1 | 10 |  | A
| TOTAL PROTEIN |  |  | g/l |  | 1 | 10 |  | A
| TOTAL PROTEIN |  |  | g/l |  | 1 | 10 |  | A
| TOTAL PROTEIN |  |  | g/l |  | 1 | 10 |  | A
| TRANSFERRIN |  |  | g/l |  | 1 | 20 |  | A
| TRANSFERRIN |  |  | g/l |  | 2 | 20 |  | A
| TRIGLYCERIDE |  |  | mmol/l |  | 2 | 25 |  | A
| TRIGLYCERIDE |  |  | mmol/l |  | 2 | 25 |  | A
| TRI-IODO THYRONINE |  |  | pmol/l |  | 1 | 11.6 |  | A | [10 empty] | Marginal failure at desirable Ricos levels due to imprecision within month (CV = 5.9%). Bias is within optimal specifications and imprecsion is within minimum specifications. No single values exceed TAE limit of 11,6%. No Qc violations noted. Calibrations performed on 11/7 which reversed bias - bias returned 26/7 (new packs not calibrated).
| TRI-IODO THYRONINE |  |  | pmol/l |  | 1 | 11.6 |  | A
| UREA |  |  | mmol/l |  | 2 | 9 |  | A
| UREA |  |  | mmol/l |  | 1 | 9 |  | A
| UREA |  |  | mmol/l |  | 1 | 9 |  | A
| UREA |  |  | mmol/l |  | 1 | 9 |  | A
| UREA URINE |  |  | mmol/l |  | 2 | 22.1 |  | C3
| UREA URINE |  |  | mmol/l |  | 2 | 22.1 |  | C3
| URIC ACID |  |  | umol/l |  | 2 | 17 |  | A
| URIC ACID |  |  | umol/l |  | 2 | 17 |  | A
| URIC ACID |  |  | umol/l |  | 2 | 17 |  | A
| URIC ACID |  |  | umol/l |  | 2 | 17 |  | A
| URIC ACID URINE |  |  | mmol/l |  | 1 | 19.4 |  | C3
| URIC ACID URINE |  |  | mmol/l |  | 1 | 19.4 |  | C3
| Urine osmolality |  |  | mosmol/kg |  | 2 | 2 |  | C4
| Urine osmolality |  |  | mosmol/kg |  | 2 | 2 |  | C4
| VITAMIN B12 II |  |  | pmol/l |  | 2 | 30 |  | A
| VITAMIN B12 II |  |  | pmol/l |  | 2 | 30 |  | A
No.Pass
[i] spreadsheet modified from Brooks, Zoe (2001). Performance-Driven Quality Control (Appendices B1 & B2) | [16 empty] | No. fail
| # Target based on |  |  |  |  | * TEa limits based on | [10 empty] | Total
| 1.inter-laboratory peer comparison |  |  |  |  | A: Westgard's recommended choice | [10 empty] | % Pass
| 2. package insert |  |  |  |  | B: Published target: NGSP or ClinChem | [10 empty] | Ave Sigma
| 3.historical/cumulative date |  |  |  |  | C: Algorithm: 1. SOTA 2. EFLM_M 3. BV_D 4. RCPA 5. Rilibak | [10 empty] | Med. Sigma
4. other ______________________

## Liaison assays
Monthly Total Error vs Total allowable error
Laboratory | GSH | [5 empty] | Department |  | C17 CHEM PATH
Date | [6 empty] | Done by | [9 empty] | Approved by
Instrument | Analyte | Control Sample | Lot number | Units | Target Value | # | TEa% | TEa (in units) | * | Mean | SD | TE (in units) | N | CV% | %Bias | SEc | Sigma | TE < TAE - Pass/ > TAE -Fail | Pathologist Comment: | QC lot usage
LIAISON XL | ALDO | ALDO L1 | 234201 | pmol/L | 215.5 | 2 | 64.2 |  | A | 208.361 | 19.74 |  | 36
LIAISON XL | ALDO | ALDO L2 | 235201 | pmol/L | 825.46 | 2 | 64.2 |  | A | 826.147 | 54.054 |  | 34
LIAISON XL | C-peptide | Liason CPEP L1 | 2124EC | ug/L | 7.43 | 2 | 20.8 |  | C | 7.138 | 0.285 |  | 52
LIAISON XL | C-peptide | Liason CPEP L3 | 2105EC | ng/ml | 2.15 | 2 | 20.8 |  | C | 2.139 | 0.073 |  | 58
LIAISON XL | LIAISON XL | Human growth hormone | 7980310 |  | 3.81 | 2 | 15 |  | C | 3.795 | 0.083 |  | 40
LIAISON XL | LIAISON XL | Human growth hormone | 9260310 |  | 17.8 | 2 | 15 |  | C | 17.276 | 0.502 |  | 41
LIAISON XL | Insulin-like growth factor I | IGF-1 L1 | 1411193 | ng/ml | 102 | 2 | 24 |  | C3 | 98.576 | 9.973 |  | 52
LIAISON XL | Insulin-like growth factor I | IGF-1 L2 | 1421193 | ng/ml | 409 | 2 | 24 |  | C3 | 380.767 | 24.504 |  | 52 | [7 empty] | All data (active lots) | All data (active lots)
LIAISON XL | Renin | RENIN L1 | 7520320 | mIU/l | 27.79 | 2 | 56.5 |  | C2 | 25.513 | 1.722 |  | 55
LIAISON XL | Renin | RENIN L2 | 7530320 | mIU/l | 116.6 | 2 | 56.5 |  | C2 | 114.467 | 5.808 |  | 52
[i] spreadsheet modified from Brooks, Zoe (2001). Performance-Driven Quality Control (Appendices B1 & B2)
| # Target based on |  |  |  |  | * TEa limits based on
| 1.inter-laboratory peer comparison |  |  |  |  | A: Westgard's recommended choice
| 2. package insert |  |  |  |  | B: Published target: NGSP or ClinChem
| 3.historical/cumulative date |  |  |  |  | C: Algorithm: 1. SOTA 2. EFLM_M 3. BV_D 4. RCPA_U 5. Thistle
4. other ______________________
| Instrument | Test
| C6000-A:E1-1 | ADRENOCORTICOTROPIC HORMONE
| C6000-A:E1-1 | ADRENOCORTICOTROPIC HORMONE
| C6000-A:E1-1 | ADRENOCORTICOTROPIC HORMONE
| C6000-A:E1-1 | ADRENOCORTICOTROPIC HORMONE
| LIAISON XL | ALDO
| LIAISON XL | ALDO
| LIAISON XL | ALDO
| LIAISON XL | ALDO
| LIAISON XL | ALDO
| C6000-A:E1-1 | ANTI-THYROID PEROXIDASE AB
| C6000-A:E1-1 | ANTI-THYROID PEROXIDASE AB
| C6000-A:E1-1 | ANTI-THYROID PEROXIDASE AB
| C6000-A:E1-1 | ANTI-THYROID PEROXIDASE AB
| C6000-A:E1-1 | ANTI-THYROID PEROXIDASE AB
| C6000-A:E1-1 | ANTI-THYROID PEROXIDASE AB
| C6000-A:E2-2 | C-peptide
| C6000-A:E2-2 | C-peptide
| LIAISON XL | C-peptide
| LIAISON XL | C-peptide
| LIAISON XL | Human growth hormone
| LIAISON XL | Human growth hormone
| LIAISON XL | Insulin-like growth factor I
| LIAISON XL | Insulin-like growth factor I
| LIAISON XL | Renin
| LIAISON XL | Renin
2021-12-10 10:03

## UOM.
Uncertainty of measurement (UOM) | [9 empty] | Optimal Bias - ratio ≤ 0.125 | [13 empty] | Optimal imprecision - ratio ≤ 0.25
[10 empty] | Desirable Bias - ratio ≤ 0.250 | [13 empty] | Desirable imprecision - ratio ≤ 0.50
[10 empty] | Minimum Bias - ratio ≤ 0.375 | [13 empty] | Minimum imprecision - ratio ≤ 0.75
[10 empty] | Bias evaluation | [13 empty] | Imprecision evaluation
Instrument | Test | QC | QC lot # | TAE target % | Obtained Tot. Error % | Sigma metric ơ | No. of data points (n) | Obtained Mean | Obtained SD | Target value | Allowable bias % | Source | % Obtained Bias | Performance Index for Bias |  | Co-efficient of variation (analytical) CVa | Co-efficient of variation (between-subject)CVb | Co-efficient of variation (within-subject)CVw | Allowable imprecision% | Source | Quality Goal Index (QGI) | Index of Individuality II | Reference Change Value (RCV) | Combined Standard Uncertainty UOM % (uc) | Extended Uncertainty UOM% (U) | Perform. Index for Imprecision (CVa/CVw)
[5 empty] | ? < TAE | World class >6 | [6 empty] | ? < Allowable bias | ≤ 0.125 | OP | ? < Allowable Imprecision | % | % |  |  | <0,8 Imprecsion |  | 95% | % |  | <0.25 | OP
[6 empty] | Excellent 5-6 | [7 empty] | >0.125 and <0.250 | DE |  | (~ CVg) | CVw |  |  | 0,8-1,2 Both |  | confidence interval |  | Coverage factor | >0,25 and <0,50 | DE
[5 empty] | Pass | Good 4-5 | [6 empty] | Pass | > 0.250 and < 0.375 | MI | Pass |  | CVi |  |  | >1,2 Inaccuracy | CVw/CVb | 1.96 |  | 2 | >0.50 and <0.75 | MI
[5 empty] | Fail | Marginal 3-4 | [6 empty] | Fail | > 0.375 |  | Fail | [9 empty] | >0.75
|  |  |  | Updated |  | Poor 2-3
[6 empty] | Unacceptable <2 | [10 empty] | https://www.westgard.com/biodatabase1.htm
Instrument | Test | QC | QC lot # | TAE target % | Obtained Tot. Error % | Sigma metric ơ | No. of data points (n) | Obtained Mean | Obtained SD | Target value | Allowable bias % | Source | % Obtained Bias | Performance Index for Bias | 0 | Co-efficient of variation (analytical) CVa | Co-efficient of variation (between-subject)CVb | Co-efficient of variation (within-subject)CVw | Allowable imprecision% | Source | Quality Goal Index (QGI) | Index of Individuality II | Reference Change Value (RCV) | Combined Standard Uncertainty UOM % (uc) | Extended Uncertainty UOM% (U) | Perform. Index for Imprecision (CVa/CVw)
| ADENOSINE DEAMINASE FLUID |  |  | 16.7% | [6 empty] | 7% | DE |  |  |  |  | 25.5% | 11.7% | 5.9% | DE | [7 empty] | 0.0223
| ADENOSINE DEAMINASE FLUID |  |  | 16.7% | [6 empty] | 7% | DE |  |  |  |  | 25.5% | 11.7% | 5.9% | DE | [7 empty] | 0.02659
| ADENOSINE DEAMINASE FLUID |  |  | 16.7% | [6 empty] | 7% | DE |  |  |  |  | 25.5% | 11.7% | 5.9% | DE | [7 empty] | 0.02585
| ADRENOCORTICOTROPIC HORMONE |  |  | 10% | [6 empty] | NO VALUE | [5 empty] | NO VALUE | 10% | NO VALUE
| ADRENOCORTICOTROPIC HORMONE |  |  | 10% | [6 empty] | NO VALUE | [5 empty] | NO VALUE | 10% | NO VALUE
| ALANINE TRANSAMINASE |  |  | 16.1% | [6 empty] | 7.7% | D |  |  |  |  | 29.3% | 10% | 5% | D | [7 empty] | 0.02415
| ALANINE TRANSAMINASE |  |  | 16.1% | [6 empty] | 7.7% | D |  |  |  |  | 29.3% | 10% | 5% | D | [7 empty] | 0.02371
| ALANINE TRANSAMINASE |  |  | 16.1% | [6 empty] | 7.7% | D |  |  |  |  | 29.3% | 10% | 5% | D | [7 empty] | 0.02009
| ALANINE TRANSAMINASE |  |  | 16.1% | [6 empty] | 7.7% | D |  |  |  |  | 29% | 10% | 5% | D | [7 empty] | 0.0272
| ALBUMIN |  |  | 10% | [6 empty] | 2.1% | MI |  |  |  |  | 4.75% | 3% | 2.4% | MI | [7 empty] | 0.02396
| ALBUMIN |  |  | 10% | [6 empty] | 2.1% | MI |  |  |  |  | 4.75% | 3% | 2.4% | MI | [7 empty] | 0.01659
| ALBUMIN |  |  | 10% | [6 empty] | 2.1% | MI |  |  |  |  | 4.75% | 3% | 2.4% | MI | [7 empty] | 0.01885
| ALBUMIN |  |  | 10% | [6 empty] | 2.1% | MI |  |  |  |  | 4.75% | 3% | 2.4% | MI | [7 empty] | 0.02898
| ALBUMIN CSF |  |  | 20% | [6 empty] | NO VALUE | [5 empty] | NO VALUE | 10% | NO VALUE | [8 empty] | 0.02431
| ALBUMIN CSF |  |  | 20% | [6 empty] | NO VALUE | [5 empty] | NO VALUE | 10% | NO VALUE | [8 empty] | 0.03067
| ALBUMIN URINE |  |  | 15% | [6 empty] | 7% | OP |  |  |  |  | 35% | 35% | 9.9% | OP | [7 empty] | 0.0132
| ALBUMIN URINE |  |  | 15% | [6 empty] | 7% | OP |  |  |  |  | 35% | 35% | 9.9% | OP | [7 empty] | 0.01444
LIAISON XL | ALDO |  |  | 36.7% | [6 empty] | 12.4% | DE |  |  |  |  | 40.1% | 29.4% | 14.7% | DE
LIAISON XL | ALDO |  |  | 36.7% | [6 empty] | 12.4% | DE |  |  |  |  | 40.1% | 29.4% | 14.7% | DE
| ALKALINE PHOSPHATASE |  |  | 14.5% | [6 empty] | 9.1% | D |  |  |  |  | 26.1% | 6.45% | 3.3% | D | [7 empty] | 0.06503
| ALKALINE PHOSPHATASE |  |  | 14.5% | [6 empty] | 9.1% | D |  |  |  |  | 26.1% | 6.45% | 3.3% | D | [7 empty] | 0.03815
| ALKALINE PHOSPHATASE |  |  | 14.5% | [6 empty] | 9.1% | D |  |  |  |  | 26.1% | 6.45% | 3.3% | D | [7 empty] | 0.03086
| ALKALINE PHOSPHATASE |  |  | 14.5% | [6 empty] | 9.1% | D |  |  |  |  | 26.1% | 6.45% | 3.3% | D | [7 empty] | 0.05579
| ALPHA 1-ANTITRYPSIN |  |  | 9.3% | [6 empty] | 4.2% | MI |  |  |  |  | 10.5% | 4.1% | 3.1% | M | [7 empty] | 0.01838
| ALPHA 1-ANTITRYPSIN |  |  | 9.3% | [6 empty] | 4.2% | MI |  |  |  |  | 10.5% | 4.1% | 3.1% | M | [7 empty] | 0.02513
| ALPHA-FETO PROTEIN |  |  | 21.9% | [6 empty] | 11.8% | DE |  |  |  |  | 45.6% | 12.2% | 6.1% | DE | [7 empty] | 0.0334
| ALPHA-FETO PROTEIN |  |  | 21.9% | [6 empty] | 11.8% | DE |  |  |  |  | 45.6% | 12.2% | 6.1% | DE | [7 empty] | 0.03181
| AMMONIA |  |  | 20% | [6 empty] | NO VALUE | [5 empty] | NO VALUE | 10% | NO VALUE | [8 empty] | 0.02586
| AMMONIA |  |  | 20% | [6 empty] | NO VALUE | [5 empty] | NO VALUE | 10% | NO VALUE | [8 empty] | 0.07796
Anti-Mullerian Hormone
Anti-Mullerian Hormone
| ANTI-THYROGLOBULIN AB |  |  | 27.6% | [6 empty] | 20.6% | DE |  |  |  |  | 82% | 8.5% | 4.3% | DE | [7 empty] | 0.07985
| ANTI-THYROGLOBULIN AB |  |  | 27.6% | [6 empty] | 20.6% | DE |  |  |  |  | 82% | 8.5% | 4.3% | DE | [7 empty] | 0.04795
| ANTI-THYROID PEROXIDASE AB |  |  | 46.2% | [6 empty] | 36.9% | DE |  |  |  |  | 147% | 11.3% | 5.7% | DE | [7 empty] | 0.07927
| ANTI-THYROID PEROXIDASE AB |  |  | 46.2% | [6 empty] | 36.9% | DE |  |  |  |  | 147% | 11.3% | 5.7% | DE | [7 empty] | 0.03987
| APOPROTEIN A |  |  | 9.1% | [6 empty] | 3.7% | DE |  |  |  |  | 13.4% | 6.5% | 3.3% | DE | [7 empty] | 0.02291
| APOPROTEIN A |  |  | 9.1% | [6 empty] | 3.7% | DE |  |  |  |  | 13.4% | 6.5% | 3.3% | DE | [7 empty] | 0.03393
| APOPROTEIN B |  |  | 11.6% | [6 empty] | 6% | DE |  |  |  |  | 22.8% | 6.9% | 3.5% | DE | [7 empty] | 0.03856
| APOPROTEIN B |  |  | 11.6% | [6 empty] | 6% | DE |  |  |  |  | 22.8% | 6.9% | 3.5% | DE | [7 empty] | 0.03877
| ASPARTATE TRANSAMINASE |  |  | 16.7% | [6 empty] | 6.54% | DE |  |  |  |  | 23.1% | 12.3% | 6.15% | DE | [7 empty] | 0.0209
| ASPARTATE TRANSAMINASE |  |  | 16.7% | [6 empty] | 6.54% | DE |  |  |  |  | 23.1% | 12.3% | 6.15% | DE | [7 empty] | 0.01834
| ASPARTATE TRANSAMINASE |  |  | 16.7% | [6 empty] | 6.54% | DE |  |  |  |  | 23.1% | 12.3% | 6.15% | DE | [7 empty] | 0.01518
| ASPARTATE TRANSAMINASE |  |  | 16.7% | [6 empty] | 6.54% | DE |  |  |  |  | 23.1% | 12.3% | 6.15% | DE | [7 empty] | 0.01863
| BETA-HCG |  |  | 15% | [6 empty] | NO VALUE | [5 empty] | NO VALUE | 15% | NO VALUE | [8 empty] | 0.02886
| BETA-HCG |  |  | 15% | [6 empty] | NO VALUE | [5 empty] | NO VALUE | 15% | NO VALUE | [8 empty] | 0.01658
| BETA-HCG |  |  | 15% | [6 empty] | NO VALUE | [5 empty] | NO VALUE | 15% | NO VALUE | [8 empty] | 0.03471
| BETA-HCG |  |  | 15% | [6 empty] | NO VALUE | [5 empty] | NO VALUE | 15% | NO VALUE | [8 empty] | 0.02347
| BICARBONATE |  |  | 17.3% | [6 empty] | 1.6% | DE |  |  |  |  | 4.8% | 4% | 2.4% | DE | [7 empty] | 0.02666
| BICARBONATE |  |  | 17.3% | [6 empty] | 1.6% | DE |  |  |  |  | 4.8% | 4% | 2.4% | DE | [7 empty] | 0.0285
| C-peptide |  |  | 20.8%
| C-peptide |  |  | 20.8%
| CA 125 |  |  | 17.7% | [6 empty] | 7.5% | OP |  |  |  |  | 54.6% | 24.7% | 6.2% | OP | [7 empty] | 0.02084
| CA 125 |  |  | 17.7% | [6 empty] | 7.5% | OP |  |  |  |  | 54.6% | 24.7% | 6.2% | OP | [7 empty] | 0.02728
| CA 19-9 |  |  | 23% | [6 empty] | 16.4% | OP |  |  |  |  | 131% | 16% | 4% | OP | [7 empty] | 0.04147
| CA 19-9 |  |  | 23% | [6 empty] | 16.4% | OP |  |  |  |  | 131% | 16% | 4% | OP | [7 empty] | 0.03948
| CALCIUM |  |  | 8.3% | [6 empty] | 1.3% | MI |  |  |  |  | 2.8% | 1.9% | 1.4% | MI | [7 empty] | 0.05371
| CALCIUM |  |  | 8.3% | [6 empty] | 1.3% | MI |  |  |  |  | 2.8% | 1.9% | 1.4% | MI | [7 empty] | 0.05597
| CALCIUM |  |  | 8.3% | [6 empty] | 1.3% | MI |  |  |  |  | 2.8% | 1.9% | 1.4% | MI | [7 empty] | 0.01282
| CALCIUM |  |  | 8.3% | [6 empty] | 1.3% | MI |  |  |  |  | 2.8% | 1.9% | 1.4% | MI | [7 empty] | 0.01437
| CALCIUM URINE |  |  | 12% | [6 empty] | 4.7% | OP |  |  |  |  | 27% | 26.2% | 6.6% | OP
| CALCIUM URINE |  |  | 12% | [6 empty] | 4.7% | OP |  |  |  |  | 27% | 26.2% | 6.6% | OP
| CARCINOEMBRYONIC AG |  |  | 12.4% | [6 empty] | 7.1% | OP |  |  |  |  | 55.6% | 12.7% | 3.2% | OP | [7 empty] | 0.01414
| CARCINOEMBRYONIC AG |  |  | 12.4% | [6 empty] | 7.1% | OP |  |  |  |  | 55.6% | 12.7% | 3.2% | OP | [7 empty] | 0.02069
| CHLORIDE |  |  | 5% | [6 empty] | 0.7% | MI |  |  |  |  | 1.5% | 1.2% | 0.9% | MI | [7 empty] | 0.01754
| CHLORIDE |  |  | 5% | [6 empty] | 0.7% | MI |  |  |  |  | 1.5% | 1.2% | 0.9% | MI | [7 empty] | 0.01613
| CHLORIDE |  |  | 5% | [6 empty] | 0.7% | MI |  |  |  |  | 1.5% | 1.2% | 0.9% | MI | [7 empty] | 0.02788
| CHLORIDE |  |  | 5% | [6 empty] | 0.7% | MI |  |  |  |  | 1.5% | 1.2% | 0.9% | MI | [7 empty] | 0.02496
| CHLORIDE URINE |  |  | 5% | [6 empty] | NO VALUE | [5 empty] | NO VALUE | 1.2% | NO VALUE | [8 empty] | 0.01763
| CHLORIDE URINE |  |  | 5% | [6 empty] | NO VALUE | [5 empty] | NO VALUE | 1.2% | NO VALUE | [8 empty] | 0.01215
| COMPLEMENT C3 |  |  | 8.4% | [6 empty] | 4.1% | DE |  |  |  |  | 15.6% | 5.2% | 2.6% | DE | [7 empty] | 0.01405
| COMPLEMENT C3 |  |  | 8.4% | [6 empty] | 4.1% | DE |  |  |  |  | 15.6% | 5.2% | 2,6% | DE | [7 empty] | 0.01724
| COMPLEMENT C4 |  |  | 8% | [6 empty] | 4.3% | OP |  |  |  |  | 33.4% | 8.9% | 4.5% | OP | [7 empty] | 0.02181
| COMPLEMENT C4 |  |  | 8% | [6 empty] | 4.3% | OP |  |  |  |  | 33.4% | 8.9% | 4.5% | OP | [7 empty] | 0.01657
| CONJUGATED BILIRUBIN |  |  | 44.5% | [6 empty] | 14.2% | DE |  |  |  |  | 43.2% | 36.8% | 18.4% | DE
| CONJUGATED BILIRUBIN |  |  | 44.5% | [6 empty] | 14.2% | DE |  |  |  |  | 43.2% | 36.8% | 18.4% | DE
| CONJUGATED BILIRUBIN |  |  | 44.5% | [6 empty] | 14.2% | DE |  |  |  |  | 43.2% | 36.8% | 18.4% | DE | [7 empty] | 0.03545
| CONJUGATED BILIRUBIN |  |  | 44.5% | [6 empty] | 14.2% | DE |  |  |  |  | 43.2% | 36.8% | 18.4% | DE | [7 empty] | 0.02572
| CONJUGATED BILIRUBIN |  |  | 44.5% | [6 empty] | 14.2% | DE |  |  |  |  | 43.2% | 36.8% | 18.4% | DE | [7 empty] | 0.03045
| CONJUGATED BILIRUBIN |  |  | 44.5% | [6 empty] | 14.2% | DE |  |  |  |  | 43.2% | 36.8% | 18.4% | DE | [7 empty] | 0.0256
| CORTISOL II |  |  | 11.4% | [6 empty] | 5.1% | OP |  |  |  |  | 38.1% | 15.2% | 3.8% | OP | [7 empty] | 0.02129
| CORTISOL II |  |  | 11.4% | [6 empty] | 5.1% | OP |  |  |  |  | 38.1% | 15.2% | 3.8% | OP | [7 empty] | 0.02862
| C-REACTIVE PROTEIN |  |  | 50.6% | [6 empty] | 22.6% | D |  |  |  |  | 83.6% | 34% | 17% | D | [7 empty] | 0.03577
| C-REACTIVE PROTEIN |  |  | 50.6% | [6 empty] | 22.6% | D |  |  |  |  | 83.6% | 34% | 17% | D | [7 empty] | 0.02662
| C-REACTIVE PROTEIN |  |  | 50.6% | [6 empty] | 22.6% | D |  |  |  |  | 83.6% | 34% | 17% | D | [7 empty] | 0.02578
| C-REACTIVE PROTEIN |  |  | 50.6% | [6 empty] | 22.6% | D |  |  |  |  | 83.6% | 34% | 17% | D | [7 empty] | 0.01752
| CREATINE KINASE Gen 2 |  |  | 24% | [6 empty] | 10.2% | D |  |  |  |  | 37.9% | 15% | 7.5% | D | [7 empty] | 0.04774
| CREATINE KINASE Gen 2 |  |  | 24% | [6 empty] | 10.2% | D |  |  |  |  | 37.9% | 15% | 7.5% | D | [7 empty] | 0.02812
| CREATININE |  |  | 15% | [6 empty] | 5.5% | MI |  |  |  |  | 14.1% | 4.4% | 3.3% | MI | [7 empty] | 0.01376
| CREATININE |  |  | 15% | [6 empty] | 5.5% | MI |  |  |  |  | 14.1% | 4.4% | 3.3% | MI | [7 empty] | 0.01488
| CREATININE |  |  | 15% | [6 empty] | 5.5% | MI |  |  |  |  | 14.1% | 4.4% | 3.3% | MI | [7 empty] | 0.02394
| CREATININE |  |  | 15% | [6 empty] | 5.5% | MI |  |  |  |  | 14.1% | 4.4% | 3.3% | MI | [7 empty] | 0.0194
| CREATININE URINE |  |  | 10% | [6 empty] | 6.4% | DE |  |  |  |  | 23% | 11% | 5.5% | DE | [7 empty] | 0.01458
| CREATININE URINE |  |  | 10% | [6 empty] | 6.4% | DE |  |  |  |  | 23% | 11% | 5.5% | DE | [7 empty] | 0.0215
| DEHYDROEPIANDROSTERONE SULPHAT |  |  | 13.08% | [6 empty] | 7.8% | DE |  |  |  |  | 30.7% | 6.35% | 3.2% | DE | [7 empty] | 0.02913
| DEHYDROEPIANDROSTERONE SULPHAT |  |  | 13.08% | [6 empty] | 7.8% | DE |  |  |  |  | 30.7% | 6.35% | 3.2% | DE | [7 empty] | 0.03052
| FERRITIN |  |  | 16.9% | [6 empty] | 5.2% | DE |  |  |  |  | 15% | 14.2% | 7.1% | DE | [7 empty] | 0.04111
| FERRITIN |  |  | 16.9% | [6 empty] | 5.2% | DE |  |  |  |  | 15% | 14.2% | 7.1% | DE | [7 empty] | 0.04757
| FOLLICLE STIMULATING HORMONE |  |  | 10.6% | [6 empty] | 6.1% | OP |  |  |  |  | 47.2% | 11% | 2.8% | OP | [7 empty] | 0.0261
| FOLLICLE STIMULATING HORMONE |  |  | 10.6% | [6 empty] | 6.1% | OP |  |  |  |  | 47.2% | 11% | 2.8% | OP | [7 empty] | 0.02407
| FREE PSA |  |  | 15% | [6 empty] | NO VALUE | [5 empty] | 46.2% | 7.1% | NO VALUE | [8 empty] | 0.02391
| FREE PSA |  |  | 15% | [6 empty] | NO VALUE | [5 empty] | 46.2% | 7.1% | NO VALUE | [8 empty] | 0.01915
| Free T4 Gen 3 |  |  | 8% | [6 empty] | 3.3% | DE |  |  |  |  | 12.1% | 5.7% | 3.6% | DE
| Free T4 Gen 3 |  |  | 8% | [6 empty] | 3.3% | DE |  |  |  |  | 12.1% | 5.7% | 3.6% | DE
| GAMMA-GLUTAMYL TRANSFERASE |  |  | 22.1% | [6 empty] | 11.06% | DE |  |  |  |  | 42.15% | 13.4% | 6.7% | DE | [7 empty] | 0.02223
| GAMMA-GLUTAMYL TRANSFERASE |  |  | 22.1% | [6 empty] | 11.06% | DE |  |  |  |  | 42.15% | 13.4% | 6.7% | DE | [7 empty] | 0.02106
| GAMMA-GLUTAMYL TRANSFERASE |  |  | 22.1% | [6 empty] | 11.06% | DE |  |  |  |  | 42.15% | 13.4% | 6.7% | DE | [7 empty] | 0.02171
| GAMMA-GLUTAMYL TRANSFERASE |  |  | 22.1% | [6 empty] | 11.06% | DE |  |  |  |  | 42.15% | 13.4% | 6.7% | DE | [7 empty] | 0.0264
| GLUCOSE |  |  | 10% | [6 empty] | 3.6% | M |  |  |  |  | 8.1% | 4.9% | 3.7% | M | [7 empty] | 0.01704
| GLUCOSE |  |  | 10% | [6 empty] | 3.6% | M |  |  |  |  | 8.1% | 4.9% | 3.7% | M | [7 empty] | 0.01512
| GLUCOSE |  |  | 10% | [6 empty] | 3.6% | M |  |  |  |  | 8.1% | 4.9% | 3.7% | M | [7 empty] | 0.02362
| GLUCOSE |  |  | 10% | [6 empty] | 3.6% | M |  |  |  |  | 8.1% | 4.9% | 3.7% | M | [7 empty] | 0.02789
| GLYCATED HAEMOGLOBIN (NGSP) |  |  | 6% | [6 empty] | 2.2% | MI |  |  |  |  | 5.7% | 1.85% | 1.4% | MI | [7 empty] | 0.0158
| GLYCATED HAEMOGLOBIN (NGSP) |  |  | 6% | [6 empty] | 2.2% | MI |  |  |  |  | 5.7% | 1.85% | 1.4% | MI
| GLYCATED HAEMOGLOBIN (NGSP) |  |  | 6% | [6 empty] | 2.2% | MI |  |  |  |  | 5.7% | 1.85% | 1.4% | MI
| GLYCATED HAEMOGLOBIN (NGSP) |  |  | 6% | [6 empty] | 2.2% | MI |  |  |  |  | 5.7% | 1.85% | 1.4% | MI | [7 empty] | 0.0139
| HAPTOGLOBIN |  |  | 13.6% | [6 empty] | 5.2% | OP |  |  |  |  | 36.4% | 20.4% | 5.1% | OP | [7 empty] | 0.03285
| HAPTOGLOBIN |  |  | 13.6% | [6 empty] | 5.2% | OP |  |  |  |  | 36.4% | 20.4% | 5.1% | OP | [7 empty] | 0.03121
| HDL CHOLESTEROL |  |  | 11.63% | [6 empty] | 5.6% | DE |  |  |  |  | 21.2% | 7.3% | 3.7% | DE | [7 empty] | 0.02019
| HDL CHOLESTEROL |  |  | 11.63% | [6 empty] | 5.6% | DE |  |  |  |  | 21.2% | 7.3% | 3.7% | DE | [7 empty] | 0.02069
| HIGH-SENSITIVITY TROPONIN T |  |  | 10% | [6 empty] | 23.8% | DE |  |  |  |  | 90% | 30.5% | 15.3% | DE | [7 empty] | 0.0211
| HIGH-SENSITIVITY TROPONIN T |  |  | 10% | [6 empty] | 23.8% | DE |  |  |  |  | 90% | 30.5% | 15.3% | DE | [7 empty] | 0.01733
| HIGH-SENSITIVITY TROPONIN T |  |  | 10% | [6 empty] | 23.8% | DE |  |  |  |  | 90% | 30.5% | 15.3% | DE | [7 empty] | 0.01218
| HIGH-SENSITIVITY TROPONIN T |  |  | 10% | [6 empty] | 23.8% | DE |  |  |  |  | 90% | 30.5% | 15.3% | DE | [7 empty] | 0.01571
LIAISON XL | hGH L1 | hGH L1 | 7980300
LIAISON XL | hGH L2 | hGH L2 | 9260300
| IGG CSF |  |  | 20% | [6 empty] | NO VALUE | [5 empty] | NO VALUE | 10% | NO VALUE | [8 empty] | 0.06034
| IGG CSF |  |  | 20% | [6 empty] | NO VALUE | [5 empty] | NO VALUE | 10% | NO VALUE | [8 empty] | 0.03942
| IMMUNOGLOBULIN A |  |  | 13.5% | [6 empty] | 9.1% | DE |  |  |  |  | 35.9% | 5.4% | 2.7% | DE | [7 empty] | 0.03323
| IMMUNOGLOBULIN A |  |  | 13.5% | [6 empty] | 9.1% | DE |  |  |  |  | 35.9% | 5.4% | 2.7% | DE | [7 empty] | 0.01928
| IMMUNOGLOBULIN G |  |  | 8% | [6 empty] | 4.3% | DE |  |  |  |  | 16.5% | 4.5% | 2.3% | DE | [7 empty] | 0.02903
| IMMUNOGLOBULIN G |  |  | 8% | [6 empty] | 4.3% | DE |  |  |  |  | 16.5% | 4.5% | 2.3% | DE | [7 empty] | 0.05216
| IMMUNOGLOBULIN M |  |  | 8.4% | [6 empty] | 6% | OP |  |  |  |  | 47.3% | 5.9% | 1.5% | OP | [7 empty] | 0.05083
| IMMUNOGLOBULIN M |  |  | 8.4% | [6 empty] | 6% | OP |  |  |  |  | 47.3% | 5.9% | 1.5% | OP | [7 empty] | 0.0227
| INORGANIC PHOSPHATE URINE |  |  | 11% | [6 empty] | 3.6% | OP |  |  |  |  | 22.6% | 18% | 4.5% | OP | [7 empty] | 0.04147
| INORGANIC PHOSPHATE URINE |  |  | 11% | [6 empty] | 3.6% | OP |  |  |  |  | 22.6% | 18% | 4.5% | OP | [7 empty] | 0.04817
| INSULIN |  |  | 25% | [6 empty] | 15.5% | DE |  |  |  |  | 58.3% | 21.1% | 10.6% | DE | [7 empty] | 0.0332
| INSULIN |  |  | 25% | [6 empty] | 15.5% | DE |  |  |  |  | 58.3% | 21.1% | 10.6% | DE | [7 empty] | 0.0283
LIAISON XL | Insulin-like growth factor I | IGF-1 L1 | 1410291 | 24 | [6 empty] | 11.9% | DE |  |  |  |  | 45.4% | 14.6% | 7.3% | DE
LIAISON XL | Insulin-like growth factor I | IGF-1 L2 | 1420291 | 24 | [6 empty] | 11.9% | DE |  |  |  |  | 45.4% | 14.6% | 7.3% | DE
| IRON |  |  | 15% | [6 empty] | 4.8% | O |  |  |  |  | 32.3% | 20.6% | 5.2% | O | [7 empty] | 0.04309
| IRON |  |  | 15% | [6 empty] | 4.8% | O |  |  |  |  | 32.3% | 20.6% | 5.2% | O | [7 empty] | 0.03441
| LACTATE |  |  | 12% | [6 empty] | 4% | OP |  |  |  |  | 16.7% | 27.2% | 6.8% | OP | [7 empty] | 0.02312
| LACTATE |  |  | 12% | [6 empty] | 4% | OP |  |  |  |  | 16.7% | 27.2% | 6.8% | OP | [7 empty] | 0.02398
| LACTATE DEHYDROGENASE |  |  | 15% | [6 empty] | 5.1% | M |  |  |  |  | 12.6% | 5.2% | 3.9% | M | [7 empty] | 0.01751
| LACTATE DEHYDROGENASE |  |  | 15% | [6 empty] | 5.1% | M |  |  |  |  | 12.6% | 5.2% | 3.9% | M | [7 empty] | 0.01517
| LACTATE DEHYDROGENASE |  |  | 15% | [6 empty] | 5.1% | M |  |  |  |  | 12.6% | 5.2% | 3.9% | M | [7 empty] | 0.02354
| LACTATE DEHYDROGENASE |  |  | 15% | [6 empty] | 5.1% | M |  |  |  |  | 12.6% | 5.2% | 3.9% | M | [7 empty] | 0.01306
| LIPASE |  |  | 37.9% | [6 empty] | 11.3% | DE |  |  |  |  | 31.8% | 32.2% | 16.1% | DE | [7 empty] | 0.10164
| LIPASE |  |  | 37.9% | [6 empty] | 11.3% | DE |  |  |  |  | 31.8% | 32.2% | 16.1% | DE | [7 empty] | 0.10947
| LITHIUM |  |  | 24.1% | [6 empty] | 6.9% | DE |  |  |  |  | 18.1% | 20.8% | 10.4% | DE | [7 empty] | 0.03556
| LITHIUM |  |  | 24.1% | [6 empty] | 6.9% | DE |  |  |  |  | 18.1% | 20.8% | 10.4% | DE | [7 empty] | 0.02969
| LUTEINISING HORMONE |  |  | 14% | [6 empty] | 4.5% | OP |  |  |  |  | 27.4% | 23% | 5.8% | OP | [7 empty] | 0.02038
| LUTEINISING HORMONE |  |  | 14% | [6 empty] | 4.5% | OP |  |  |  |  | 27.4% | 23% | 5.8% | OP | [7 empty] | 0.01493
| MAGNESIUM |  |  | 11% | [6 empty] | 2.8% | MI |  |  |  |  | 6.4% | 3.6% | 2.7% | MI | [7 empty] | 0.02213
| MAGNESIUM |  |  | 11% | [6 empty] | 2.8% | MI |  |  |  |  | 6.4% | 3.6% | 2.7% | MI | [7 empty] | 0.02359
| MAGNESIUM |  |  | 11% | [6 empty] | 2.8% | MI |  |  |  |  | 6.4% | 3.6% | 2.7% | MI | [7 empty] | 0.02788
| MAGNESIUM |  |  | 11% | [6 empty] | 2.8% | MI |  |  |  |  | 6.4% | 3.6% | 2.7% | MI | [7 empty] | 0.03012
| MAGNESIUM URINE |  |  | 25% | [6 empty] | 2.8% | MI |  |  |  |  | 6.4% | 3.6% | 2.7% | MI | [7 empty] | 0.03992
| MAGNESIUM URINE |  |  | 25% | [6 empty] | 2.8% | MI |  |  |  |  | 6.4% | 3.6% | 2.7% | MI | [7 empty] | 0.04288
| OESTRADIOL III |  |  | 26.86% | [6 empty] | 8.3% | DE |  |  |  |  | 24.4% | 22.5% | 11.25% | DE | [7 empty] | 0.02801
| OESTRADIOL III |  |  | 26.86% | [6 empty] | 8.3% | DE |  |  |  |  | 24.4% | 22.5% | 11.25% | DE | [7 empty] | 0.02563
| Osmolality |  |  | 1.5% | [6 empty] | 0.4% | DE |  |  |  |  | 1.2% | 1.3% | 0.7% | DE | [7 empty] | 0.01591
| Osmolality |  |  | 1.5% | [6 empty] | 0.4% | DE |  |  |  |  | 1.2% | 1.3% | 0.7% | DE | [7 empty] | 0.01304
| PARACETAMOL 2 |  |  | 10%
| PARACETAMOL 2 |  |  | 10%
| PARATHYROID HORMONE |  |  | 15.1% | [6 empty] | 4.4% | OP |  |  |  |  | 23.8% | 25.9% | 6.5% | OP | [7 empty] | 0.02102
| PARATHYROID HORMONE |  |  | 15.1% | [6 empty] | 4.4% | OP |  |  |  |  | 23.8% | 25.9% | 6.5% | OP | [7 empty] | 0.01516
| PHOSPHATE |  |  | 10.1% | [6 empty] | 3.38% | DE |  |  |  |  | 10.8% | 8.2% | 4.08% | DE | [7 empty] | 0.01644
| PHOSPHATE |  |  | 10.1% | [6 empty] | 3.38% | DE |  |  |  |  | 10.8% | 8.2% | 4.08% | DE | [7 empty] | 0.01215
| PHOSPHATE |  |  | 10.1% | [6 empty] | 3.38% | DE |  |  |  |  | 10.8% | 8.2% | 4.08% | DE | [7 empty] | 0.0182
| PHOSPHATE |  |  | 10.1% | [6 empty] | 3.38% | DE |  |  |  |  | 10.8% | 8.2% | 4.08% | DE | [7 empty] | 0.02421
| POTASSIUM |  |  | 5.61% | [6 empty] | 1.81% | OP |  |  |  |  | 5.6% | 4.6% | 2.3% | OP | [7 empty] | 0.03059
| POTASSIUM |  |  | 5.61% | [6 empty] | 1.81% | OP |  |  |  |  | 5.6% | 4.6% | 2.3% | OP | [7 empty] | 0.01372
| POTASSIUM |  |  | 5.61% | [6 empty] | 1.8% | DE |  |  |  |  | 5.6% | 4.6% | 2.3% | DE | [7 empty] | 0.05402
| POTASSIUM |  |  | 5.6% | [6 empty] | 1.8% | DE |  |  |  |  | 5.6% | 4.6% | 2.3% | DE | [7 empty] | 0.02748
| POTASSIUM URINE |  |  | 10% | [6 empty] | 1.8% | DE |  |  |  |  | 5.6% | 4.6% | 2.3% | DE | [7 empty] | 0.04593
| POTASSIUM URINE |  |  | 10% | [6 empty] | 1.8% | DE |  |  |  |  | 5.6% | 4.6% | 2.3% | DE | [7 empty] | 0.02468
| proBNP |  |  | 13% | [6 empty] | 4.7% | DE |  |  |  |  | 16% | 10% | 5% | DE | [7 empty] | 0.09689
| proBNP |  |  | 13% | [6 empty] | 4.7% | DE |  |  |  |  | 22.2% | 10% | 5% | DE | [7 empty] | 0.098
| PROCALCITONIN (SENSITIVE) |  |  | 20%
| PROCALCITONIN (SENSITIVE) |  |  | 20%
| PROGESTERONE III |  |  | 25% | [6 empty] | NO VALUE | [5 empty] | NO VALUE | NO VALUE | NO VALUE | [8 empty] | 0.01489
| PROGESTERONE III |  |  | 25% | [6 empty] | NO VALUE | [5 empty] | NO VALUE | NO VALUE | NO VALUE | [8 empty] | 0.0146
| PROLACTIN |  |  | 14.7% | [6 empty] | 5.2% | OP |  |  |  |  | 34.95% | 23% | 5.8% | [8 empty] | 0.01847
| PROLACTIN |  |  | 14.7% | [6 empty] | 5.2% | OP |  |  |  |  | 34.95% | 23% | 8.8% | [8 empty] | 0.02297
| PROSTATE-SPECIFIC AG |  |  | 16.8% | [6 empty] | 9.3% | OP |  |  |  |  | 72.4% | 18.1% | 4.5% | OP | [7 empty] | 0.01574
| PROSTATE-SPECIFIC AG |  |  | 16.8% | [6 empty] | 9.3% | OP |  |  |  |  | 72.4% | 18.1% | 4.5% | OP | [7 empty] | 0.01324
| PROTEIN CSF |  |  | 10% | [6 empty] | 9.3% | OP |  |  |  |  | 72.4% | 18.1% | 4.5% | OP | [7 empty] | 0.01312
| PROTEIN CSF |  |  | 10% | [6 empty] | 9.3% | OP |  |  |  |  | 72.4% | 18.1% | 4.5% | OP | [7 empty] | 0.01891
| PROTEIN URINE |  |  | 20% | [6 empty] | NO VALUE | [5 empty] | NO VALUE | 10% | NO VALUE | [8 empty] | 0.01763
| PROTEIN URINE |  |  | 20% | [6 empty] | NO VALUE | [5 empty] | NO VALUE | 10% | NO VALUE | [8 empty] | 0.02509
LIAISON XL | Renin | RENIN L1 | 7520310 | 20
LIAISON XL | Renin | RENIN L2 | 7530310 | 20
| RHD3 |  |  | 6% | [6 empty] | 2.2% | MI |  |  |  |  | 5.7% | 1.85% | 8.9 | OP
| RHD3 |  |  | 6% | [6 empty] | 2.2% | MI |  |  |  |  | 5.7% | 1.85% | 8.9 | OP
| RHEUMATOID FACTOR |  |  | 13.5% | [6 empty] | 6.5% | DE |  |  |  |  | 24.5% | 8.5% | 4.3% | DE | [7 empty] | 0.04347
| RHEUMATOID FACTOR |  |  | 13.5% | [6 empty] | 6.5% | DE |  |  |  |  | 24.5% | 8.5% | 4.3% | DE | [7 empty] | 0.03452
| SERUM FOLATE III |  |  | 39% | [6 empty] | 19.2% | DE |  |  |  |  | 73% | 24% | 12% | DE
| SERUM FOLATE III |  |  | 39% | [6 empty] | 19.2% | DE |  |  |  |  | 73% | 24% | 12% | DE
| SEX HORMONE BINDING GLOBULIN |  |  | 20.42% | [6 empty] | 9.66% | DE |  |  |  |  | 36.35% | 13.05% | 6.53% | DE | [7 empty] | 0.0294
| SEX HORMONE BINDING GLOBULIN |  |  | 20.42% | [6 empty] | 9.66% | DE |  |  |  |  | 36.35% | 13.05% | 6.53% | DE | [7 empty] | 0.02744
| SODIUM |  |  | 2% | [6 empty] | 0.3% | MI |  |  |  |  | 0.7% | 0.6% | 0.5% | MI | [7 empty] | 0.0376
| SODIUM |  |  | 2% | [6 empty] | 0.3% | MI |  |  |  |  | 0.7% | 0.6% | 0.5% | MI | [7 empty] | 0.03156
| SODIUM |  |  | 2% | [6 empty] | 0.3% | MI |  |  |  |  | 0.7% | 0.6% | 0.5% | MI | [7 empty] | 0.02981
| SODIUM |  |  | 2% | [6 empty] | 0.3% | MI |  |  |  |  | 0.7% | 0.6% | 0.5% | MI | [7 empty] | 0.00936
| SODIUM URINE |  |  | 10% | [6 empty] | 0.3% | MI |  |  |  |  | 0.7% | 0.6% | 0.5% | MI | [7 empty] | 0.01233
| SODIUM URINE |  |  | 10% | [6 empty] | 0.3% | MI |  |  |  |  | 0.7% | 0.6% | 0.5% | MI | [7 empty] | 0.0222
| TESTOSTERONE |  |  | 13.61% | [6 empty] | 5.98% | DE |  |  |  |  | 22.05% | 9.25% | 4.63% | DE | [7 empty] | 0.0105
| TESTOSTERONE |  |  | 13.6% | [6 empty] | 5.98% | DE |  |  |  |  | 22.05% | 9.25% | 4.63% | DE | [7 empty] | 0.00912
| THYROGLOBULIN v2 |  |  | 11% | [6 empty] | 5.2% | OP |  |  |  |  | 39% | 14% | 3.5% | OP | [7 empty] | 0.02512
| THYROGLOBULIN v2 |  |  | 11% | [6 empty] | 5.2% | OP |  |  |  |  | 39% | 14% | 3.5% | OP | [7 empty] | 0.03862
| THYROID STIMULATING HORMONE |  |  | 11.9% | [6 empty] | 3.9% | OP |  |  |  |  | 24.6% | 19.3% | 4.8% | OP | [7 empty] | 0.03734
| THYROID STIMULATING HORMONE |  |  | 11.9% | [6 empty] | 3.9% | OP |  |  |  |  | 24.6% | 19.3% | 4.8% | OP | [7 empty] | 0.06565
| THYROID STIMULATING HORMONE |  |  | 11.9% | [6 empty] | 3.9% | OP |  |  |  |  | 24.6% | 19.3% | 4.8% | OP | [7 empty] | 0.03457
| THYROID STIMULATING HORMONE |  |  | 11.9% | [6 empty] | 3.9% | OP |  |  |  |  | 24.6% | 19.3% | 4.8% | OP | [7 empty] | 0.01472
| TOTAL BILIRUBIN |  |  | 20% | [6 empty] | 8.3% | D |  |  |  |  | 26.6% | 20% | 10% | D | [7 empty] | 0.01274
| TOTAL BILIRUBIN |  |  | 20% | [6 empty] | 8.3% | D |  |  |  |  | 26.6% | 20% | 10% | D | [7 empty] | 0.0157
| TOTAL BILIRUBIN |  |  | 20% | [6 empty] | 8.3% | D |  |  |  |  | 26.6% | 20% | 10% | D
| TOTAL BILIRUBIN |  |  | 20% | [6 empty] | 8.3% | D |  |  |  |  | 26.6% | 20% | 10% | D
| TOTAL BILIRUBIN |  |  | 20% | [6 empty] | 8.3% | D |  |  |  |  | 26% | 20% | 10% | D | [7 empty] | 0.00999
| TOTAL BILIRUBIN |  |  | 20% | [6 empty] | 8.3% | D |  |  |  |  | 26% | 20% | 10% | D | [7 empty] | 0.03736
| TOTAL CHOLESTEROL |  |  | 10% | [6 empty] | 4.3% | D |  |  |  |  | 16.2% | 5.3% | 2.7% | D | [7 empty] | 0.04302
| TOTAL CHOLESTEROL |  |  | 10% | [6 empty] | 4.3% | D |  |  |  |  | 16.2% | 5.3% | 2.7% | D | [7 empty] | 0.03495
| TOTAL CHOLESTEROL |  |  | 10% | [6 empty] | 4.3% | D |  |  |  |  | 16.2% | 5.3% | 2.7% | D | [7 empty] | 0.03007
| TOTAL CHOLESTEROL |  |  | 10% | [6 empty] | 4.3% | D |  |  |  |  | 16.2% | 5.3% | 2.7% | D | [7 empty] | 0.02853
| TOTAL PROTEIN |  |  | 10% | [6 empty] | 2% | MI |  |  |  |  | 4.7% | 2.75% | 2.1% | MI | [7 empty] | 0.03508
| TOTAL PROTEIN |  |  | 10% | [6 empty] | 2% | MI |  |  |  |  | 4.7% | 2.75% | 2.1% | MI | [7 empty] | 0.02726
| TOTAL PROTEIN |  |  | 10% | [6 empty] | 2% | MI |  |  |  |  | 4.7% | 2.75% | 2.1% | MI | [7 empty] | 0.02651
| TOTAL PROTEIN |  |  | 10% | [6 empty] | 2% | MI |  |  |  |  | 4.7% | 2.75% | 2.1% | MI | [7 empty] | 0.02955
| TRANSFERRIN |  |  | 20% | [6 empty] | 5.4% | M |  |  |  |  | 13.9% | 3.8% | 2.9% | M | [7 empty] | 0.01965
| TRANSFERRIN |  |  | 20% | [6 empty] | 5.4% | M |  |  |  |  | 13.9% | 3.8% | 2.9% | M | [7 empty] | 0.02688
| TRIGLYCERIDE |  |  | 25% | [6 empty] | 10.5% | D |  |  |  |  | 37% | 20% | 10% | D | [7 empty] | 0.03625
| TRIGLYCERIDE |  |  | 25% | [6 empty] | 10.5% | D |  |  |  |  | 37% | 20% | 10% | D | [7 empty] | 0.01461
| TRI-IODO THYRONINE |  |  | 11.3% | [6 empty] | 4.8% | DE |  |  |  |  | 17.6% | 7.9% | 4% | DE | [7 empty] | 0.02021
| TRI-IODO THYRONINE |  |  | 11.3% | [6 empty] | 4.8% | DE |  |  |  |  | 17.6% | 7.9% | 4% | DE | [7 empty] | 0.01414
| UREA |  |  | 9% | [6 empty] | 3.1% | O |  |  |  |  | 20.9% | 13.9% | 3.5% | O | [7 empty] | 0.017
| UREA |  |  | 9% | [6 empty] | 3.1% | O |  |  |  |  | 20.9% | 13.9% | 3.5% | O | [7 empty] | 0.01792
| UREA |  |  | 9% | [6 empty] | 3.1% | O |  |  |  |  | 20.9% | 13.9% | 3.5% | O | [7 empty] | 0.01987
| UREA |  |  | 9% | [6 empty] | 3.1% | O |  |  |  |  | 20.9% | 13.9% | 3.5% | O | [7 empty] | 0.02889
| UREA URINE |  |  | 11% | [6 empty] | 2.8% | OP |  |  |  |  | 18.7% | 12.1% | 3% | OP | [7 empty] | 0.02567
| UREA URINE |  |  | 11% | [6 empty] | 2.8% | OP |  |  |  |  | 18.7% | 12.1% | 3% | OP | [7 empty] | 0.01543
| URIC ACID |  |  | 17% | [6 empty] | 8.9% | M |  |  |  |  | 22.4% | 8.2% | 6.2% | M | [7 empty] | 0.01739
| URIC ACID |  |  | 17% | [6 empty] | 8.9% | M |  |  |  |  | 22.4% | 8.2% | 6.2% | M | [7 empty] | 0.02623
| URIC ACID |  |  | 17% | [6 empty] | 8.9% | M |  |  |  |  | 22.4% | 8.2% | 6.2% | M | [7 empty] | 0.02003
| URIC ACID |  |  | 17% | [6 empty] | 8.9% | M |  |  |  |  | 22.4% | 8.2% | 6.2% | M | [7 empty] | 0.03824
| URIC ACID URINE |  |  | 9.7% | [6 empty] | 4.9% | DE |  |  |  |  | 17.5% | 8.6% | 4.3% | DE | [7 empty] | 0.01965
| URIC ACID URINE |  |  | 9.7% | [6 empty] | 4.9% | DE |  |  |  |  | 17.5% | 8.6% | 4.3% | DE | [7 empty] | 0.02354
| Urine osmolality |  |  | 2% | [6 empty] | 8.1% | OP |  |  |  |  | 57.9% | 28.3% | 17.1% | OP | [7 empty] | 0.04285
| Urine osmolality |  |  | 2% | [6 empty] | 8.1% | OP |  |  |  |  | 57.9% | 28.3% | 17.1% | OP | [7 empty] | 0.0213
| VITAMIN B12 II |  |  | 20% | [6 empty] | 17.7% | DE |  |  |  |  | 69% | 15% | 7.5% | DE
| VITAMIN B12 II |  |  | 20% | [6 empty] | 17.7% | DE |  |  |  |  | 69% | 15% | 7.5% | DE
0.01779
0.01892
0.01816
0.0172
[12 empty] | # Target based on Westgard | [7 empty] | # Target based on Westgard
[12 empty] | OP - optimal specifications -Ricos | [7 empty] | OP. optimal specifications
[12 empty] | DE - desirable specifications -Ricos | [7 empty] | DE. desirable specifications
[12 empty] | MI - minimum specifications - Ricos | [7 empty] | MI. minimum specifications
| Prep. by | E. Samsodien | [9 empty] | M -EFLM - Minimum specs | [7 empty] | M -EFLM - Minimum specs
[12 empty] | D - EFLM - Desirable specs | [7 empty] | D - EFLM - Desirable specs
| Date: | [10 empty] | O - EFLM - Optimal specs | [7 empty] | O - EFLM - Optimal specs
Reviewed by :
Date :
Note on A1: nhls: based on Westgard desirable performance specifications
Note on K5: nhls: Source may be package insert, peer or cumulative mean
Note on L5: nhls: https://www.westgard.com/biodatabase1.htm
Note on N5: nhls: Is obtained bias less than allowable bias at desirable specifications?
Note on O5: nhls: White GH, Farrance I. Uncertainty of measurement in quantitative medical testing. A laboratory implementation guide. Clin Biochem Rev 2004; 25: Suppl S: 1–24.
Note on Q5: nhls: analytical CV from lab data
Note on R5: nhls: between subject biological variation
Note on S5: nhls; within subject biological variation
Note on T5: nhls: https://www.westgard.com/biodatabase1.htm
Note on W5: nhls: indicates the practical benefit of the reference range for the analyte II > 1.4 - indicates greater practical benefit of the normal range
Note on X5: Reference change value (RCV) for detectinga clinically significant change in two consecutive results for the same analytic for the same patient.
Note on Y5: nhls: combined uncertainties of measurement - square root of the individual squares of the analytical CV and within-subject biological variation CV
Note on Z5: nhls: uc X coverage factor
Note on AA5: : Index = CVi/CVw Source: . Fraser CG, Hyltoft Petersen P, Libeer JC, Ricos C. Pro posals for setting generally applicable quality goals solely based on biology. Ann Clin Biochem 1997; 34: 8–12.

## Raw Data
Instrument | Test | QC | QC lot No. | Target mean | Target SD | SDI | Number of values | Bias | Calculated mean | Calculated SD | CV% | Usage
C6000-B:C2 | ADENOSINE DEAMINASE FLUID | ADA 1 | 101220 | 29,800 | 1,800 | 0,382 | 23 | 0,023 | 30,487 | 0,597 | 1,959 | In use
C6000-B:C2 | ADENOSINE DEAMINASE FLUID | ADA 2 | 201220 | 141,000 | 8,000 | 0,306 | 23 | 0,017 | 143,448 | 7,007 | 4,885 | In use
C6000-B:C2 | ADENOSINE DEAMINASE FLUID | ADA 0 | 01220 | 11,300 | 0,600 | 0,217 | 23 | 0,012 | 11,430 | 0,365 | 3,192 | In use
C6000-A:E1-1 | ADRENOCORTICOTROPIC HORMONE | PC MM1 | 563238 | 9,500 | 0,500 | 0,500 | 5 | 0,026 | 9,750 | 0,327 | 3,357 | In use
C6000-A:E1-1 | ADRENOCORTICOTROPIC HORMONE | PC MM2 | 563239 | 207,300 | 10,800 | 0,037 | 5 | 0,002 | 207,700 | 8,132 | 3,915 | In use
C6000-A:C1 | ALANINE TRANSAMINASE | PCCC 1 | 525205 | 48,900 | 2,900 | -0,584 | 51 | -0,035 | 47,206 | 1,214 | 2,573 | In use
C6000-A:C1 | ALANINE TRANSAMINASE | PCCC 2 | 535719 | 127,000 | 8,000 | -0,780 | 47 | -0,049 | 120,760 | 3,924 | 3,249 | In use
C6000-B:C2 | ALANINE TRANSAMINASE | PCCC 2-B | 535719 | 127,000 | 8,000 | -0,632 | 41 | -0,040 | 121,941 | 4,897 | 4,016 | In use
C6000-B:C2 | ALANINE TRANSAMINASE | PCCC 1-B | 525205 | 48,900 | 2,900 | -0,480 | 47 | -0,028 | 47,509 | 1,637 | 3,445 | In use
C6000-B:C2 | ALBUMIN | s-PCCC 2-B | 535719 | 48,800 | 2,900 | -0,133 | 67 | -0,008 | 48,415 | 1,533 | 3,166 | In use
C6000-A:C1 | ALBUMIN | PCCC 1 | 525205 | 32,100 | 1,900 | 0,217 | 44 | 0,013 | 32,511 | 1,153 | 3,545 | In use
C6000-A:C1 | ALBUMIN | PCCC 2 | 535719 | 48,800 | 2,900 | 0,073 | 43 | 0,004 | 49,012 | 1,589 | 3,242 | In use
C6000-B:C2 | ALBUMIN | s-PCCC 1-B | 525205 | 32,100 | 1,900 | -0,256 | 68 | -0,015 | 31,613 | 1,013 | 3,204 | In use
C6000-A:C1 | ALBUMIN CSF | PPPUC TPC3 | 646927 | 116,000 | 9,000 | 0,787 | 47 | 0,061 | 123,081 | 5,580 | 4,534 | In use
C6000-A:C1 | ALBUMIN CSF | LIQ SFC2 | 56532 | 318,100 | 13,400 | -0,939 | 45 | -0,040 | 305,522 | 9,848 | 3,223 | In use
C6000-A:C1 | ALBUMIN URINE | Ran-Liq 3 | 1238UC | 62,400 | 6,250 | 0,036 | 35 | 0,004 | 62,623 | 1,446 | 2,310 | In use
C6000-A:C1 | ALBUMIN URINE | Ran-Liq 2 | 1237UC | 24,000 | 2,400 | 0,096 | 38 | 0,010 | 24,232 | 1,417 | 5,848 | In use
C6000-B:C2 | ALKALINE PHOSPHATASE | s-PCCC 2-B | 535719 | 267,000 | 16,000 | -0,762 | 70 | -0,046 | 254,814 | 7,024 | 2,757 | In use
C6000-A:C1 | ALKALINE PHOSPHATASE | s-PCCC 2 | 535719 | 267,000 | 16,000 | -0,581 | 70 | -0,035 | 257,700 | 9,177 | 3,561 | In use
C6000-A:C1 | ALKALINE PHOSPHATASE | s-PCCC 1 | 525205 | 101,000 | 6,000 | -0,734 | 74 | -0,044 | 96,595 | 4,614 | 4,776 | In use
C6000-B:C2 | ALKALINE PHOSPHATASE | s-PCCC 1-B | 525205 | 101,000 | 6,000 | -0,876 | 70 | -0,052 | 95,743 | 3,829 | 3,999 | In use
C6000-B:C2 | ALPHA 1-ANTITRYPSIN | PCCC 2-B | 535719 | 1,320 | 0,110 | 0,296 | 29 | 0,025 | 1,353 | 0,019 | 1,420 | In use
C6000-B:C2 | ALPHA 1-ANTITRYPSIN | PCCC 1-B | 525205 | 0,883 | 0,071 | 0,458 | 35 | 0,037 | 0,916 | 0,015 | 1,639 | In use
C6000-A:E1-1 | ALPHA-FETO PROTEIN | RIAP L3 | 2181EC | 218,000 | 22,000 | -0,369 | 37 | -0,037 | 209,876 | 7,808 | 3,720 | In use
C6000-A:E1-1 | ALPHA-FETO PROTEIN | RIAP L1 | 2179EC | 10,600 | 1,040 | -0,357 | 36 | -0,035 | 10,229 | 0,575 | 5,620 | In use
C6000-B:C2 | AMMONIA | AMM-N | 644004 | 62,500 | 4,200 | 0,254 | 34 | 0,017 | 63,566 | 1,733 | 2,726 | In use
C6000-B:C2 | AMMONIA | AMM-P | 644006 | 228,000 | 15,000 | 0,029 | 35 | 0,002 | 228,442 | 3,354 | 1,468 | In use
C6000-A:E2-2 | Anti-Mullerian Hormone | PC AMH 1 | 670151 | 8,000 | 0,560 | -0,190 | 6 | -0,013 | 7,893 | 0,207 | 2,624 | In use
C6000-A:E2-2 | Anti-Mullerian Hormone | PC AMH 2 | 670154 | 39,700 | 1,990 | -0,104 | 6 | -0,005 | 39,493 | 1,468 | 3,717 | In use
C6000-A:E2-1 | ANTI-THYROGLOBULIN AB | PC THYRO1 | 669238 | 58,000 | 5,800 | 0,759 | 28 | 0,076 | 62,404 | 2,520 | 4,038 | In use
C6000-A:E2-1 | ANTI-THYROGLOBULIN AB | PC THYRO2 | 669239 | 195,000 | 19,500 | 0,402 | 28 | 0,040 | 202,836 | 9,586 | 4,726 | In use
C6000-A:E1-1 | ANTI-THYROID PEROXIDASE AB | PC THYRO1 | 669238 | 47,300 | 4,730 | 0,568 | 9 | 0,057 | 49,987 | 4,824 | 9,652 | In use
C6000-A:E1-1 | ANTI-THYROID PEROXIDASE AB | PC THYRO2 | 669239 | 109,000 | 7,630 | 0,740 | 9 | 0,052 | 114,644 | 5,843 | 5,097 | In use
C6000-B:C2 | APOPROTEIN A | PCCC 2-B | 535719 | 1,660 | 0,130 | 0,358 | 10 | 0,028 | 1,707 | 0,058 | 3,399 | In use
C6000-B:C2 | APOPROTEIN A | PCCC 1-B | 525205 | 1,100 | 0,090 | 0,319 | 10 | 0,026 | 1,129 | 0,044 | 3,894 | In use
C6000-B:C2 | APOPROTEIN B | PCCC 2-B | 535719 | 0,762 | 0,061 | 0,220 | 9 | 0,018 | 0,775 | 0,012 | 1,546 | In use
C6000-B:C2 | APOPROTEIN B | PCCC 1-B | 525205 | 0,484 | 0,039 | 0,815 | 9 | 0,066 | 0,516 | 0,013 | 2,540 | In use
C6000-A:C1 | ASPARTATE TRANSAMINASE | PCCC 1 | 525205 | 44,700 | 2,700 | -0,415 | 43 | -0,025 | 43,579 | 0,786 | 1,804 | In use
C6000-A:C1 | ASPARTATE TRANSAMINASE | PCCC 2 | 535719 | 143,000 | 9,000 | -0,690 | 42 | -0,043 | 136,790 | 2,683 | 1,961 | In use
C6000-B:C2 | ASPARTATE TRANSAMINASE | PCCC 2-B | 535719 | 143,000 | 9,000 | -0,623 | 41 | -0,039 | 137,390 | 2,727 | 1,985 | In use
C6000-B:C2 | ASPARTATE TRANSAMINASE | PCCC 1-B | 525205 | 44,700 | 2,700 | -0,342 | 48 | -0,021 | 43,777 | 1,293 | 2,954 | In use
C6000-B:E3-2 | BETA-HCG | RIAP L1-B | 2179EC | 11,400 | 1,140 | 0,155 | 37 | 0,015 | 11,576 | 0,269 | 2,327 | In use
C6000-B:E3-2 | BETA-HCG | RIAP L3-B | 2181EC | 272,000 | 27,000 | -0,359 | 36 | -0,036 | 262,297 | 7,018 | 2,675 | In use
C6000-A:E1-1 | BETA-HCG | RIAP L3 | 2181EC | 272,000 | 27,000 | -0,587 | 38 | -0,058 | 256,147 | 9,869 | 3,853 | In use
C6000-A:E1-1 | BETA-HCG | RIAP L1 | 2179EC | 11,400 | 1,140 | 0,035 | 38 | 0,003 | 11,440 | 0,620 | 5,422 | In use
C6000-B:C2 | BICARBONATE | AMM-N | 644004 | 17,000 | 1,000 | 0,812 | 34 | 0,048 | 17,812 | 0,766 | 4,302 | In use
C6000-B:C2 | BICARBONATE | AMM-P | 644006 | 30,500 | 1,800 | -0,441 | 33 | -0,026 | 29,706 | 0,658 | 2,214 | In use
C6000-A:E2-1 | CA 125 | RIAP L3 | 2181EC | 182,000 | 18,000 | 0,506 | 39 | 0,050 | 191,110 | 6,437 | 3,368 | In use
C6000-A:E2-1 | CA 125 | RIAP L1 | 2179EC | 19,000 | 1,900 | 0,377 | 37 | 0,038 | 19,717 | 0,909 | 4,613 | In use
C6000-A:E2-1 | CA 19-9 | RIAP L3 | 2181EC | 165,000 | 16,500 | 0,255 | 40 | 0,025 | 169,203 | 7,397 | 4,371 | In use
C6000-A:E2-1 | CA 19-9 | RIAP L1 | 2179EC | 20,800 | 2,100 | 0,259 | 39 | 0,026 | 21,344 | 0,947 | 4,436 | In use
C6000-A:C1 | CALCIUM | PCCC 1 | 525205 | 2,140 | 0,090 | -0,543 | 44 | -0,023 | 2,091 | 0,033 | 1,595 | In use
C6000-A:C1 | CALCIUM | PCCC 2 | 535719 | 3,470 | 0,140 | -0,478 | 43 | -0,019 | 3,403 | 0,050 | 1,470 | In use
C6000-B:C2 | CALCIUM | PCCC 2-B | 535719 | 3,470 | 0,140 | -0,364 | 40 | -0,015 | 3,419 | 0,043 | 1,265 | In use
C6000-B:C2 | CALCIUM | PCCC 1-B | 525205 | 2,140 | 0,090 | -0,644 | 48 | -0,027 | 2,082 | 0,051 | 2,429 | In use
C6000-B:C2 | CALCIUM URINE | Ran-Liq 3 | 1238UC | 2,780 | 0,140 | 0,020 | 28 | 0,001 | 2,783 | 0,061 | 2,202 | In use
C6000-B:C2 | CALCIUM URINE | Ran-Liq 2 | 1237UC | 1,870 | 0,100 | 0,125 | 28 | 0,007 | 1,883 | 0,045 | 2,387 | In use
C6000-A:E1-2 | CARCINOEMBRYONIC AG | RIAP L3 | 2181EC | 50,000 | 5,000 | 0,566 | 34 | 0,057 | 52,831 | 2,601 | 4,923 | In use
C6000-A:E1-2 | CARCINOEMBRYONIC AG | RIAP L1 | 2179EC | 4,080 | 0,410 | 0,405 | 32 | 0,041 | 4,246 | 0,229 | 5,400 | In use
C6000-B:ISE1 | CHLORIDE | s-PCCC 2-B | 535719 | 106,000 | 3,000 | 0,223 | 82 | 0,006 | 106,668 | 2,191 | 2,054 | In use
C6000-A:ISE1 | CHLORIDE | s-PCCC 2 | 535719 | 103,000 | 3,000 | 0,802 | 95 | 0,023 | 105,407 | 1,817 | 1,724 | In use
C6000-A:ISE1 | CHLORIDE | s-PCCC 1 | 525205 | 83,700 | 2,500 | 0,308 | 95 | 0,009 | 84,471 | 1,672 | 1,979 | In use
C6000-B:ISE1 | CHLORIDE | s-PCCC 1-B | 525205 | 83,700 | 2,500 | 0,664 | 83 | 0,020 | 85,359 | 2,038 | 2,388 | In use
C6000-B:ISE1 | CHLORIDE URINE | Ran-Liq 3 | 1238UC | 181,000 | 13,500 | -0,146 | 33 | -0,011 | 179,024 | 2,437 | 1,361 | In use
C6000-B:ISE1 | CHLORIDE URINE | Ran-Liq 2 | 1237UC | 76,200 | 5,700 | 0,567 | 32 | 0,042 | 79,431 | 1,791 | 2,255 | In use
C6000-B:C2 | COMPLEMENT C3 | PCCC 2-B | 535719 | 1,550 | 0,120 | 0,438 | 28 | 0,034 | 1,603 | 0,055 | 3,416 | In use
C6000-B:C2 | COMPLEMENT C3 | PCCC 1-B | 525205 | 0,926 | 0,074 | 0,718 | 34 | 0,057 | 0,979 | 0,033 | 3,336 | In use
C6000-B:C2 | COMPLEMENT C4 | PCCC 2-B | 535719 | 0,227 | 0,018 | -0,200 | 27 | -0,016 | 0,223 | 0,003 | 1,444 | In use
C6000-B:C2 | COMPLEMENT C4 | PCCC 1-B | 525205 | 0,145 | 0,012 | -0,490 | 32 | -0,041 | 0,139 | 0,003 | 2,256 | In use
C6000-A:C1 | CONJUGATED BILIRUBIN | MAS BILIRUBIN 3 | BC24033A | 177,000 | 17,300 | -0,130 | 36 | -0,013 | 174,744 | 5,192 | 2,971 | In use
C6000-B:C2 | CONJUGATED BILIRUBIN | MAS BILIRUBIN-B | BC24033A | 177,000 | 17,300 | -0,080 | 35 | -0,008 | 175,609 | 6,322 | 3,600 | In use
C6000-A:C1 | CONJUGATED BILIRUBIN | PCCC 1 | 525205 | 16,400 | 1,300 | -0,132 | 45 | -0,010 | 16,229 | 0,607 | 3,738 | In use
C6000-A:C1 | CONJUGATED BILIRUBIN | PCCC 2 | 535719 | 44,900 | 3,600 | -0,343 | 44 | -0,027 | 43,666 | 1,324 | 3,032 | In use
C6000-B:C2 | CONJUGATED BILIRUBIN | PCCC 2-B | 535719 | 44,900 | 3,600 | -0,306 | 37 | -0,025 | 43,797 | 1,305 | 2,979 | In use
C6000-B:C2 | CONJUGATED BILIRUBIN | PCCC 1-B | 525205 | 16,400 | 1,300 | -0,021 | 44 | -0,002 | 16,373 | 0,505 | 3,082 | In use
C6000-A:E1-2 | CORTISOL II | RIAP L1 | 2179EC | 201,000 | 25,000 | -0,119 | 39 | -0,015 | 198,015 | 6,715 | 3,391 | In use
C6000-A:E1-2 | CORTISOL II | RIAP L2 | 2228EC | 660,000 | 82,500 | -0,009 | 38 | -0,001 | 659,237 | 19,443 | 2,949 | In use
C6000-A:E2-2 | C-peptide | RIAP L1 | 2179EC | 1,940 | 0,160 | -1,167 | 54 | -0,096 | 1,753 | 0,125 | 7,149 | In use
C6000-A:E2-2 | C-peptide | RIAP L2 | 2228EC | 6,160 | 0,770 | -0,301 | 52 | -0,038 | 5,928 | 0,406 | 6,854 | In use
C6000-A:C1 | C-REACTIVE PROTEIN | PCCC 1 | 525205 | 5,680 | 0,210 | -0,654 | 41 | -0,024 | 5,543 | 0,101 | 1,815 | In use
C6000-A:C1 | C-REACTIVE PROTEIN | PCCC 2 | 535719 | 51,400 | 3,400 | -0,455 | 40 | -0,030 | 49,853 | 1,486 | 2,981 | In use
C6000-B:C2 | C-REACTIVE PROTEIN | PCCC 2-B | 535719 | 51,400 | 3,400 | -0,128 | 38 | -0,008 | 50,964 | 1,570 | 3,080 | In use
C6000-B:C2 | C-REACTIVE PROTEIN | PCCC 1-B | 525205 | 5,680 | 0,210 | -0,700 | 44 | -0,026 | 5,533 | 0,125 | 2,251 | In use
C6000-A:C1 | CREATINE KINASE Gen 2 | PCCC 1 | 525205 | 153,000 | 9,000 | -0,393 | 39 | -0,023 | 149,462 | 1,945 | 1,301 | In use
C6000-A:C1 | CREATINE KINASE Gen 2 | PCCC 2 | 535719 | 267,000 | 16,000 | -0,263 | 38 | -0,016 | 262,789 | 2,961 | 1,127 | In use
C6000-A:C1 | CREATININE | PCCC 1 | 525205 | 88,700 | 5,300 | 0,068 | 117 | 0,004 | 89,060 | 1,984 | 2,228 | In use
C6000-A:C1 | CREATININE | PCCC 2 | 535719 | 329,000 | 20,000 | 0,199 | 115 | 0,012 | 332,983 | 7,183 | 2,157 | In use
C6000-B:C2 | CREATININE | PCCC 2-B | 535719 | 329,000 | 20,000 | -0,424 | 109 | -0,026 | 320,514 | 6,271 | 1,957 | In use
C6000-B:C2 | CREATININE | PCCC 1-B | 525205 | 88,700 | 5,300 | -0,379 | 111 | -0,023 | 86,694 | 1,589 | 1,832 | In use
C6000-A:C1 | CREATININE URINE | Ran-Liq 3 | 1238UC | 15,500 | 1,550 | -0,290 | 92 | -0,029 | 15,050 | 0,428 | 2,841 | In use
C6000-A:C1 | CREATININE URINE | Ran-Liq 2 | 1237UC | 6,680 | 0,670 | -0,326 | 96 | -0,033 | 6,461 | 0,171 | 2,652 | In use
C6000-A:E2-2 | DEHYDROEPIANDROSTERONE SULPHAT | RIAP L1 | 2179EC | 2,680 | 0,270 | 0,311 | 36 | 0,031 | 2,764 | 0,276 | 9,993 | In use
C6000-A:E2-2 | DEHYDROEPIANDROSTERONE SULPHAT | RIAP L2 | 2228EC | 17,500 | 1,750 | 0,233 | 38 | 0,023 | 17,908 | 1,611 | 8,998 | In use
C6000-A | FERRITIN | PC V1 | 612148 | 146,000 | 13,100 | 0,170 | 3 | 0,015 | 148,233 | 4,136 | 2,790 | In use
C6000-A:E2-1 | FERRITIN | PC V1 | 612148 | 146,000 | 13,100 | 0,032 | 51 | 0,003 | 146,425 | 2,376 | 1,623 | In use
C6000-A:E2-1 | FERRITIN | sRIAP L1 | 2179EC | 20,570 | 2,030 | 1,559 | 92 | 0,154 | 23,735 | 1,108 | 4,669 | In use
C6000-A:E2-1 | FERRITIN | RIAP L3 | 2181EC | 318,000 | 32,000 | 0,802 | 45 | 0,081 | 343,664 | 10,233 | 2,978 | In use
C6000-A:E2-1 | FERRITIN | sRIAP L3 | 2181EC | 318,000 | 32,000 | 0,765 | 90 | 0,077 | 342,476 | 7,125 | 2,081 | In use
C6000-A:E1-2 | FOLLICLE STIMULATING HORMONE | RIAP L3 | 2181EC | 54,200 | 5,400 | 0,484 | 37 | 0,048 | 56,815 | 1,230 | 2,165 | In use
C6000-A:E1-2 | FOLLICLE STIMULATING HORMONE | RIAP L1 | 2179EC | 6,030 | 0,605 | 0,927 | 39 | 0,093 | 6,591 | 0,155 | 2,359 | In use
C6000-A:E1-2 | FREE PSA | RIAP L1 | 2179EC | 1,190 | 0,149 | -0,294 | 35 | -0,037 | 1,146 | 0,048 | 4,206 | In use
C6000-A:E1-2 | FREE PSA | RIAP L2 | 2228EC | 9,210 | 1,150 | -0,141 | 34 | -0,018 | 9,047 | 0,255 | 2,819 | In use
C6000-A:E1-1 | Free T4 Gen 3 | sRIAP L1 | 2179EC | 13,700 | 1,700 | 0,147 | 77 | 0,018 | 13,951 | 0,712 | 5,101 | In use
C6000-A:E1-1 | Free T4 Gen 3 | sRIAP L2 | 2228EC | 29,900 | 3,750 | 0,099 | 76 | 0,012 | 30,272 | 1,288 | 4,253 | In use
C6000-A:C1 | GAMMA-GLUTAMYL TRANSFERASE | PCCC 1 | 525205 | 58,200 | 3,500 | -0,258 | 44 | -0,016 | 57,295 | 0,851 | 1,486 | In use
C6000-A:C1 | GAMMA-GLUTAMYL TRANSFERASE | PCCC 2 | 535719 | 229,000 | 14,000 | -0,350 | 43 | -0,021 | 224,093 | 3,315 | 1,479 | In use
C6000-B:C2 | GAMMA-GLUTAMYL TRANSFERASE | PCCC 2-B | 535719 | 229,000 | 14,000 | -0,474 | 38 | -0,029 | 222,368 | 2,432 | 1,094 | In use
C6000-B:C2 | GAMMA-GLUTAMYL TRANSFERASE | PCCC 1-B | 525205 | 58,200 | 3,500 | -0,518 | 44 | -0,031 | 56,386 | 1,298 | 2,302 | In use
C6000-A:C1 | GLUCOSE | PCCC 1 | 525205 | 5,760 | 0,290 | -0,409 | 41 | -0,021 | 5,641 | 0,081 | 1,432 | In use
C6000-A:C1 | GLUCOSE | PCCC 2 | 535719 | 13,600 | 0,700 | -0,656 | 39 | -0,034 | 13,141 | 0,170 | 1,294 | In use
C6000-B:C2 | GLUCOSE | PCCC 2-B | 535719 | 13,600 | 0,700 | -0,504 | 35 | -0,026 | 13,247 | 0,268 | 2,020 | In use
C6000-B:C2 | GLUCOSE | PCCC 1-B | 525205 | 5,760 | 0,290 | -0,327 | 41 | -0,016 | 5,665 | 0,116 | 2,054 | In use
C6000-A:C1 | GLYCATED HAEMOGLOBIN (NGSP) | PCA1P | 617429 | 9,690 | 0,580 | -0,350 | 55 | -0,021 | 9,487 | 0,206 | 2,166 | In use
C6000-B:C2 | GLYCATED HAEMOGLOBIN (NGSP) | PCA1P | 617429 | 9,690 | 0,580 | 0,078 | 81 | 0,005 | 9,735 | 0,132 | 1,355 | In use
C6000-B:C2 | GLYCATED HAEMOGLOBIN (NGSP) | PCA1N | 617428 | 5,700 | 0,340 | 0,150 | 81 | 0,009 | 5,751 | 0,105 | 1,824 | In use
C6000-A:C1 | GLYCATED HAEMOGLOBIN (NGSP) | PCA1N-VER | 617428 | 5,700 | 0,340 | -0,128 | 55 | -0,008 | 5,656 | 0,096 | 1,693 | In use
C6000-B:C2 | HAPTOGLOBIN | PCCC 2-B | 535719 | 1,210 | 0,100 | 0,451 | 40 | 0,037 | 1,255 | 0,016 | 1,296 | In use
C6000-B:C2 | HAPTOGLOBIN | PCCC 1-B | 525205 | 0,765 | 0,061 | 0,661 | 44 | 0,053 | 0,805 | 0,014 | 1,702 | In use
C6000-A:C1 | HDL CHOLESTEROL | PCCC 1 | 525205 | 0,719 | 0,058 | 1,056 | 39 | 0,085 | 0,780 | 0,018 | 2,343 | In use
C6000-A:C1 | HDL CHOLESTEROL | PCCC 2 | 535719 | 1,520 | 0,120 | 0,772 | 38 | 0,061 | 1,613 | 0,040 | 2,496 | In use
C6000-B:E3-1 | HIGH-SENSITIVITY TROPONIN T | BIO CTC1-B | 89041 | 17,990 | 0,560 | 1,123 | 39 | 0,035 | 18,619 | 0,435 | 2,338 | In use
C6000-A:E2-1 | HIGH-SENSITIVITY TROPONIN T | BIO CTC1 | 89041 | 17,990 | 0,560 | 1,095 | 41 | 0,034 | 18,603 | 1,153 | 6,199 | In use
C6000-A:E2-1 | HIGH-SENSITIVITY TROPONIN T | PC TN1 | 657678 | 29,100 | 2,620 | 0,166 | 23 | 0,015 | 29,536 | 1,852 | 6,270 | In use
C6000-B:E3-1 | HIGH-SENSITIVITY TROPONIN T | PC TN1-B | 657678 | 29,100 | 2,620 | 0,236 | 22 | 0,021 | 29,719 | 0,449 | 1,511 | In use
C6000-A:C1 | IGG CSF | PNPUC TPC3 | 673292 | 21,200 | 1,700 | -0,753 | 40 | -0,060 | 19,920 | 1,074 | 5,392 | In use
C6000-A:C1 | IGG CSF | PPPUC TPC3 | 646927 | 129,000 | 10,000 | -0,264 | 40 | -0,020 | 126,363 | 3,356 | 2,656 | In use
C6000-B:C2 | IMMUNOGLOBULIN A | PCCC 2-B | 535719 | 2,260 | 0,180 | -0,092 | 26 | -0,007 | 2,243 | 0,035 | 1,559 | In use
C6000-B:C2 | IMMUNOGLOBULIN A | PCCC 1-B | 525205 | 1,540 | 0,120 | -0,476 | 31 | -0,037 | 1,483 | 0,028 | 1,861 | In use
C6000-B:C2 | IMMUNOGLOBULIN G | PCCC 2-B | 535719 | 11,200 | 0,700 | 0,744 | 28 | 0,047 | 11,721 | 0,285 | 2,433 | In use
C6000-B:C2 | IMMUNOGLOBULIN G | PCCC 1-B | 525205 | 7,350 | 0,440 | 0,758 | 33 | 0,045 | 7,684 | 0,201 | 2,622 | In use
C6000-B:C2 | IMMUNOGLOBULIN M | PCCC 2-B | 535719 | 1,020 | 0,080 | 0,456 | 26 | 0,036 | 1,056 | 0,019 | 1,812 | In use
C6000-B:C2 | IMMUNOGLOBULIN M | PCCC 1-B | 525205 | 0,623 | 0,050 | 1,043 | 32 | 0,084 | 0,675 | 0,019 | 2,831 | In use
C6000-B:C2 | INORGANIC PHOSPHATE URINE | Ran-Liq 3 | 1238UC | 13,800 | 1,400 | 0,539 | 26 | 0,055 | 14,555 | 0,451 | 3,102 | In use
C6000-B:C2 | INORGANIC PHOSPHATE URINE | Ran-Liq 2 | 1237UC | 8,230 | 0,830 | 0,545 | 26 | 0,055 | 8,682 | 0,267 | 3,076 | In use
C6000-A:E1-2 | INSULIN | RIAP L3 | 2181EC | 28,800 | 3,600 | -0,092 | 45 | -0,012 | 28,468 | 1,004 | 3,529 | In use
C6000-A:E1-2 | INSULIN | RIAP L2 | 2228EC | 16,100 | 2,000 | -0,148 | 41 | -0,018 | 15,803 | 0,875 | 5,536 | In use
C6000-A:C1 | IRON | PCCC 1 | 525205 | 19,300 | 1,200 | 0,165 | 43 | 0,010 | 19,498 | 0,646 | 3,315 | In use
C6000-A:C1 | IRON | PCCC 2 | 535719 | 44,200 | 2,700 | -0,241 | 41 | -0,015 | 43,549 | 0,683 | 1,568 | In use
C6000-B:C2 | LACTATE | PCCC 2-B | 535719 | 3,820 | 0,230 | -0,081 | 39 | -0,005 | 3,801 | 0,066 | 1,723 | In use
C6000-B:C2 | LACTATE | PCCC 1-B | 525205 | 1,780 | 0,110 | -0,305 | 45 | -0,019 | 1,746 | 0,032 | 1,828 | In use
C6000-A:C1 | LACTATE DEHYDROGENASE | PCCC 1 | 525205 | 164,000 | 10,000 | 0,493 | 44 | 0,030 | 168,932 | 4,008 | 2,373 | In use
C6000-A:C1 | LACTATE DEHYDROGENASE | PCCC 2 | 535719 | 299,000 | 18,000 | 0,397 | 43 | 0,024 | 306,140 | 9,528 | 3,112 | In use
C6000-B:C2 | LACTATE DEHYDROGENASE | PCCC 2-B | 535719 | 299,000 | 18,000 | 0,133 | 36 | 0,008 | 301,389 | 8,846 | 2,935 | In use
C6000-B:C2 | LACTATE DEHYDROGENASE | PCCC 1-B | 525205 | 164,000 | 10,000 | 0,407 | 42 | 0,025 | 168,071 | 6,482 | 3,857 | In use
C6000-A:C1 | LIPASE | PCCC 1 | 525205 | 45,000 | 2,700 | -0,460 | 48 | -0,028 | 43,758 | 1,659 | 3,791 | In use
C6000-A:C1 | LIPASE | PCCC 2 | 535719 | 100,000 | 6,000 | -0,517 | 47 | -0,031 | 96,900 | 3,516 | 3,628 | In use
C6000-B:C2 | LIPOPROTEIN A2 | PC LPA 1 | 520719 | 44,800 | 3,000 | -0,251 | 8 | -0,017 | 44,048 | 1,033 | 2,345 | In use
C6000-B:C2 | LIPOPROTEIN A2 | PC LPA 2 | 520720 | 112,000 | 7,000 | 0,546 | 8 | 0,034 | 115,819 | 3,991 | 3,446 | In use
C6000-B:C2 | LITHIUM | PCCC 2-B | 535719 | 1,810 | 0,070 | -0,113 | 43 | -0,004 | 1,802 | 0,070 | 3,885 | In use
C6000-B:C2 | LITHIUM | PCCC 1-B | 525205 | 0,935 | 0,037 | -0,196 | 49 | -0,008 | 0,928 | 0,044 | 4,715 | In use
C6000-A:E1-2 | LUTEINISING HORMONE | RIAP L3 | 2181EC | 55,020 | 1,960 | 1,289 | 41 | 0,046 | 57,546 | 1,251 | 2,174 | In use
C6000-A:E1-2 | LUTEINISING HORMONE | RIAP L1 | 2179EC | 4,260 | 0,327 | -0,819 | 39 | -0,063 | 3,992 | 0,124 | 3,103 | In use
C6000-A:C1 | MAGNESIUM | PCCC 1 | 525205 | 0,793 | 0,032 | 0,221 | 48 | 0,009 | 0,800 | 0,025 | 3,104 | In use
C6000-A:C1 | MAGNESIUM | PCCC 2 | 535719 | 1,360 | 0,050 | 0,044 | 45 | 0,002 | 1,362 | 0,045 | 3,282 | In use
C6000-B:C2 | MAGNESIUM | PCCC 2-B | 535719 | 1,360 | 0,050 | -0,151 | 47 | -0,006 | 1,352 | 0,064 | 4,735 | In use
C6000-B:C2 | MAGNESIUM | PCCC 1-B | 525205 | 0,793 | 0,032 | -0,145 | 53 | -0,006 | 0,788 | 0,042 | 5,371 | In use
C6000-B:C2 | MAGNESIUM URINE | Ran-Liq 3 | 1238UC | 6,170 | 0,620 | 0,269 | 31 | 0,027 | 6,337 | 0,257 | 4,056 | In use
C6000-B:C2 | MAGNESIUM URINE | Ran-Liq 2 | 1237UC | 2,880 | 0,290 | 0,481 | 31 | 0,048 | 3,019 | 0,159 | 5,273 | In use
C6000-A:E1-1 | OESTRADIOL III | RIAP L3 | 2181EC | 1480,000 | 148,000 | 0,318 | 36 | 0,032 | 1527,083 | 28,983 | 1,898 | In use
C6000-A:E1-1 | OESTRADIOL III | RIAP L1 | 2179EC | 198,000 | 20,000 | 0,647 | 38 | 0,065 | 210,939 | 9,922 | 4,704 | In use
OSMO1 | Osmolality | Serim BIOLIQ 2 | 45952 | 398,000 | 9,300 | 0,154 | 30 | 0,004 | 399,433 | 3,213 | 0,804 | In use
OSMO1 | Osmolality | Serim BIOLIQ 1 | 45931 | 311,000 | 3,833 | 0,362 | 31 | 0,004 | 312,387 | 2,591 | 0,829 | In use
C6000-B:C2 | PARACETAMOL 2 | RIAP L1-B | 2179EC | 153,000 | 15,500 | -0,305 | 34 | -0,031 | 148,271 | 6,781 | 4,573 | In use
C6000-B:C2 | PARACETAMOL 2 | RIAP L2-B | 2228EC | 323,000 | 21,700 | -0,441 | 34 | -0,030 | 313,426 | 16,837 | 5,372 | In use
C6000-A:E2-1 | PARATHYROID HORMONE | PC V1 | 612148 | 6,230 | 0,560 | 0,048 | 45 | 0,004 | 6,257 | 0,349 | 5,577 | In use
C6000-A:E2-1 | PARATHYROID HORMONE | PC V2 | 595765 | 19,100 | 1,530 | 1,427 | 24 | 0,114 | 21,284 | 0,950 | 4,462 | Study
C6000-A:C1 | PHOSPHATE | PCCC 1 | 525205 | 1,190 | 0,060 | -0,543 | 47 | -0,027 | 1,157 | 0,020 | 1,723 | In use
C6000-A:C1 | PHOSPHATE | PCCC 2 | 535719 | 2,760 | 0,140 | -0,597 | 47 | -0,030 | 2,676 | 0,040 | 1,487 | In use
C6000-B:C2 | PHOSPHATE | PCCC 2-B | 535719 | 2,760 | 0,140 | -0,247 | 37 | -0,013 | 2,725 | 0,089 | 3,268 | In use
C6000-B:C2 | PHOSPHATE | PCCC 1-B | 525205 | 1,190 | 0,060 | -0,076 | 44 | -0,004 | 1,185 | 0,034 | 2,879 | In use
C6000-B:ISE1 | POTASSIUM | s-PCCC 2-B | 535719 | 7,260 | 0,220 | 0,214 | 85 | 0,006 | 7,307 | 0,095 | 1,295 | In use
C6000-A:ISE1 | POTASSIUM | s-PCCC 2 | 535719 | 7,260 | 0,220 | 0,525 | 93 | 0,016 | 7,375 | 0,128 | 1,732 | In use
C6000-A:ISE1 | POTASSIUM | s-PCCC 1 | 525205 | 3,490 | 0,100 | 0,946 | 94 | 0,027 | 3,585 | 0,075 | 2,094 | In use
C6000-B:ISE1 | POTASSIUM | s-PCCC 1-B | 525205 | 3,490 | 0,100 | 0,489 | 84 | 0,014 | 3,539 | 0,077 | 2,164 | In use
C6000-B:ISE1 | POTASSIUM URINE | Ran-Liq 3 | 1238UC | 72,380 | 1,480 | 0,476 | 31 | 0,010 | 73,085 | 1,737 | 2,377 | In use
C6000-B:ISE1 | POTASSIUM URINE | Ran-Liq 2 | 1237UC | 27,500 | 2,050 | 0,779 | 32 | 0,058 | 29,098 | 0,661 | 2,270 | In use
C6000-A:E1-1 | proBNP | PC CARD1 | 629283 | 132,000 | 7,920 | -0,172 | 32 | -0,010 | 130,638 | 3,150 | 2,412 | In use
C6000-A:E1-1 | proBNP | PC CARD2 | 629284 | 4930,000 | 296,000 | -0,384 | 32 | -0,023 | 4816,375 | 133,986 | 2,782 | In use
C6000-A:E2-1 | PROCALCITONIN (SENSITIVE) | PC PCT1 | 671597 | 0,490 | 0,030 | 0,374 | 35 | 0,023 | 0,501 | 0,013 | 2,638 | In use
C6000-A:E2-1 | PROCALCITONIN (SENSITIVE) | PC PCT2 | 671599 | 9,390 | 0,560 | -0,340 | 35 | -0,020 | 9,199 | 0,235 | 2,550 | In use
C6000-A:E2-2 | PROGESTERONE III | RIAP L3 | 2181EC | 98,300 | 9,750 | 0,453 | 36 | 0,045 | 102,719 | 5,543 | 5,397 | In use
C6000-A:E2-2 | PROGESTERONE III | RIAP L1 | 2179EC | 3,490 | 0,350 | 0,697 | 37 | 0,070 | 3,734 | 0,387 | 10,377 | In use
C6000-A:E1-2 | PROLACTIN | RIAP L3 | 2181EC | 73,132 | 7,300 | -0,435 | 36 | -0,043 | 69,954 | 2,859 | 4,087 | In use
C6000-A:E1-2 | PROLACTIN | RIAP L2 | 2228EC | 33,041 | 3,300 | -0,579 | 34 | -0,058 | 31,129 | 1,331 | 4,277 | In use
C6000-A:E2-2 | PROSTATE-SPECIFIC AG | RIAP L1 | 2179EC | 1,630 | 0,205 | -0,052 | 46 | -0,007 | 1,619 | 0,049 | 3,015 | In use
C6000-A:E2-2 | PROSTATE-SPECIFIC AG | RIAP L2 | 2228EC | 12,200 | 1,525 | -0,010 | 46 | -0,001 | 12,185 | 0,616 | 5,059 | In use
C6000-A:C1 | PROTEIN CSF | PNPUC TPC3 | 673292 | 0,218 | 0,022 | -0,148 | 47 | -0,015 | 0,215 | 0,006 | 2,651 | In use
C6000-A:C1 | PROTEIN CSF | PPPUC TPC3 | 646927 | 1,550 | 0,160 | -0,150 | 47 | -0,015 | 1,526 | 0,031 | 2,004 | In use
C6000-A:C1 | PROTEIN URINE | Ran-Liq 3 | 1238UC | 0,710 | 0,070 | -0,120 | 34 | -0,012 | 0,702 | 0,020 | 2,831 | In use
C6000-A:C1 | PROTEIN URINE | Ran-Liq 2 | 1237UC | 0,130 | 0,010 | 0,255 | 33 | 0,020 | 0,133 | 0,005 | 3,411 | In use
C6000-B:C2 | RHD3 | PPA1c H | 617429 | 9,820 | 0,590 | -0,522 | 75 | -0,031 | 9,512 | 0,195 | 2,047 | In use
C6000-B:C2 | RHD3 | PNA1c H | 617428 | 5,700 | 0,340 | 0,085 | 75 | 0,005 | 5,729 | 0,100 | 1,753 | In use
C6000-B:C2 | RHEUMATOID FACTOR | RFCO1 | 661133 | 22,700 | 1,500 | 0,835 | 27 | 0,055 | 23,952 | 0,687 | 2,870 | In use
C6000-B:C2 | RHEUMATOID FACTOR | RFCO2 | 661134 | 51,800 | 2,600 | 0,640 | 27 | 0,032 | 53,463 | 1,318 | 2,466 | In use
C6000-A:E2-2 | SERUM FOLATE III | RIAP L3 | 2181EC | 21,400 | 2,550 | 0,761 | 49 | 0,091 | 23,341 | 1,552 | 6,651 | In use
C6000-A:E2-2 | SERUM FOLATE III | RIAP L2 | 2228EC | 12,027 | 0,715 | 0,037 | 54 | 0,002 | 12,054 | 0,937 | 7,773 | In use
C6000-A:E2-2 | SEX HORMONE BINDING GLOBULIN | RIAP L1 | 2179EC | 9,700 | 0,970 | 0,373 | 34 | 0,037 | 10,062 | 0,340 | 3,380 | In use
C6000-A:E2-2 | SEX HORMONE BINDING GLOBULIN | RIAP L2 | 2228EC | 34,100 | 3,400 | 0,400 | 38 | 0,040 | 35,459 | 1,136 | 3,203 | In use
C6000-B:ISE1 | SODIUM | s-PCCC 2-B | 535719 | 135,000 | 4,000 | 0,506 | 84 | 0,015 | 137,024 | 1,598 | 1,166 | In use
C6000-A:ISE1 | SODIUM | s-PCCC 2 | 535719 | 135,000 | 4,000 | 0,842 | 95 | 0,025 | 138,368 | 1,935 | 1,399 | In use
C6000-A:ISE1 | SODIUM | s-PCCC 1 | 525205 | 112,000 | 3,000 | 0,811 | 95 | 0,022 | 114,432 | 1,849 | 1,616 | In use
C6000-B:ISE1 | SODIUM | s-PCCC 1-B | 525205 | 112,000 | 3,000 | 0,433 | 84 | 0,012 | 113,298 | 1,789 | 1,579 | In use
C6000-B:ISE1 | SODIUM URINE | Ran-Liq 3 | 1238UC | 173,000 | 10,500 | -0,130 | 33 | -0,008 | 171,636 | 1,917 | 1,117 | In use
C6000-B:ISE1 | SODIUM URINE | Ran-Liq 2 | 1237UC | 86,800 | 5,200 | -0,406 | 32 | -0,024 | 84,688 | 2,867 | 3,386 | In use
C6000-A:E2-1 | TESTOSTERONE | RIAP L3 | 2181EC | 28,700 | 2,850 | -0,280 | 37 | -0,028 | 27,903 | 1,277 | 4,577 | In use
C6000-A:E2-1 | TESTOSTERONE | RIAP L2 | 2228EC | 14,700 | 1,450 | 0,292 | 35 | 0,029 | 15,123 | 0,541 | 3,576 | In use
C6000-A:E1-1 | THYROGLOBULIN v2 | RIAP L3 | 2181EC | 60,000 | 7,500 | 0,738 | 35 | 0,092 | 65,533 | 1,617 | 2,468 | In use
C6000-A:E1-1 | THYROGLOBULIN v2 | RIAP L1 | 2179EC | 1,890 | 0,235 | 0,725 | 36 | 0,090 | 2,060 | 0,074 | 3,601 | In use
C6000-B:E3-1 | THYROID STIMULATING HORMONE | RIAP L1-B | 2179EC | 0,136 | 0,014 | -0,370 | 41 | -0,037 | 0,131 | 0,007 | 5,723 | In use
C6000-B:E3-1 | THYROID STIMULATING HORMONE | RIAP L2-B | 2228EC | 2,540 | 0,255 | 0,150 | 41 | 0,015 | 2,578 | 0,087 | 3,361 | In use
C6000-A:E2-2 | THYROID STIMULATING HORMONE | RIAP L1 | 2179EC | 0,136 | 0,014 | -0,684 | 59 | -0,068 | 0,127 | 0,007 | 5,680 | In use
C6000-A:E2-2 | THYROID STIMULATING HORMONE | RIAP L2 | 2228EC | 2,540 | 0,255 | 0,018 | 59 | 0,002 | 2,545 | 0,093 | 3,664 | In use
C6000-A:C1 | TOTAL BILIRUBIN | MAS BILIRUBIN 3 | BC24033A | 340,000 | 22,700 | 0,330 | 43 | 0,022 | 347,493 | 12,628 | 3,634 | In use
C6000-B:C2 | TOTAL BILIRUBIN | MAS BILIRUBIN-B | BC24033A | 340,000 | 22,700 | 0,098 | 36 | 0,007 | 342,214 | 11,835 | 3,458 | In use
C6000-A:C1 | TOTAL BILIRUBIN | PCCC 1 | 525205 | 16,400 | 1,000 | 0,243 | 53 | 0,015 | 16,643 | 0,762 | 4,578 | In use
C6000-A:C1 | TOTAL BILIRUBIN | PCCC 2 | 535719 | 62,300 | 3,700 | 0,174 | 51 | 0,010 | 62,943 | 2,669 | 4,240 | In use
C6000-B:C2 | TOTAL BILIRUBIN | PCCC 2-B | 535719 | 62,300 | 3,700 | -0,217 | 39 | -0,013 | 61,497 | 2,604 | 4,234 | In use
C6000-B:C2 | TOTAL BILIRUBIN | PCCC 1-B | 525205 | 16,400 | 1,000 | -0,036 | 45 | -0,002 | 16,364 | 0,692 | 4,226 | In use
C6000-B:C2 | TOTAL CHOLESTEROL | s-PCCC 2-B | 535719 | 4,400 | 0,220 | -0,579 | 77 | -0,029 | 4,273 | 0,069 | 1,626 | In use
C6000-A:C1 | TOTAL CHOLESTEROL | s-PCCC 2 | 535719 | 4,400 | 0,220 | 0,154 | 76 | 0,008 | 4,434 | 0,120 | 2,714 | In use
C6000-A:C1 | TOTAL CHOLESTEROL | s-PCCC 1 | 525205 | 2,420 | 0,120 | 0,338 | 78 | 0,017 | 2,461 | 0,075 | 3,053 | In use
C6000-B:C2 | TOTAL CHOLESTEROL | s-PCCC 1-B | 525205 | 2,420 | 0,120 | -0,287 | 78 | -0,014 | 2,386 | 0,041 | 1,699 | In use
C6000-A:C1 | TOTAL PROTEIN | PCCC 1 | 525205 | 48,200 | 1,900 | -0,810 | 52 | -0,032 | 46,662 | 1,376 | 2,949 | In use
C6000-A:C1 | TOTAL PROTEIN | PCCC 2 | 535719 | 76,600 | 3,100 | -0,900 | 50 | -0,036 | 73,810 | 2,299 | 3,115 | In use
C6000-B:C2 | TOTAL PROTEIN | PCCC 2-B | 535719 | 76,600 | 3,100 | -0,282 | 39 | -0,011 | 75,726 | 2,177 | 2,875 | In use
C6000-B:C2 | TOTAL PROTEIN | PCCC 1-B | 525205 | 48,200 | 1,900 | -0,240 | 46 | -0,009 | 47,743 | 1,453 | 3,043 | In use
C6000-A:C1 | TRANSFERRIN | PCCC 1 | 525205 | 1,890 | 0,110 | 0,200 | 54 | 0,012 | 1,912 | 0,041 | 2,147 | In use
C6000-A:C1 | TRANSFERRIN | PCCC 2 | 535719 | 3,330 | 0,200 | -0,056 | 53 | -0,003 | 3,319 | 0,077 | 2,331 | In use
C6000-A:C1 | TRIGLYCERIDE | PCCC 1 | 525205 | 1,340 | 0,070 | -0,201 | 49 | -0,011 | 1,326 | 0,019 | 1,444 | In use
C6000-A:C1 | TRIGLYCERIDE | PCCC 2 | 535719 | 2,460 | 0,120 | -0,326 | 47 | -0,016 | 2,421 | 0,031 | 1,286 | In use
C6000-A:E1-2 | TRI-IODO THYRONINE | RIAP L1 | 2179EC | 4,960 | 0,620 | 0,008 | 38 | 0,001 | 4,965 | 0,158 | 3,181 | In use
C6000-A:E1-2 | TRI-IODO THYRONINE | RIAP L2 | 2228EC | 15,600 | 1,950 | -0,152 | 37 | -0,019 | 15,304 | 0,513 | 3,350 | In use
C6000-B:C2 | UREA | s-PCCC 2-B | 535719 | 20,600 | 1,000 | -0,646 | 80 | -0,031 | 19,954 | 0,495 | 2,481 | In use
C6000-A:C1 | UREA | s-PCCC 2 | 535719 | 20,600 | 1,000 | -0,734 | 82 | -0,036 | 19,866 | 0,940 | 4,733 | In use
C6000-A:C1 | UREA | s-PCCC 1 | 525205 | 6,760 | 0,340 | -0,520 | 83 | -0,026 | 6,583 | 0,317 | 4,814 | In use
C6000-B:C2 | UREA | s-PCCC 1-B | 525205 | 6,760 | 0,340 | -0,398 | 81 | -0,020 | 6,625 | 0,165 | 2,498 | In use
C6000-B:C2 | UREA URINE | Ran-Liq 3 | 1238UC | 289,000 | 29,000 | -0,369 | 28 | -0,037 | 278,286 | 8,199 | 2,946 | In use
C6000-A:C1 | UREA URINE | Ran-Liq 3 | 1238UC | 289,000 | 29,000 | -0,007 | 20 | -0,001 | 288,800 | 18,071 | 6,257 | In use
C6000-B:C2 | UREA URINE | Ran-Liq 2 | 1237UC | 149,000 | 15,000 | -0,218 | 28 | -0,022 | 145,725 | 4,704 | 3,228 | In use
C6000-A:C1 | UREA URINE | Ran-Liq 2 | 1237UC | 149,000 | 15,000 | 0,161 | 23 | 0,016 | 151,413 | 7,622 | 5,034 | In use
C6000-A:C1 | URIC ACID | PCCC 1 | 525205 | 0,282 | 0,014 | 0,207 | 41 | 0,010 | 0,285 | 0,005 | 1,895 | In use
C6000-A:C1 | URIC ACID | PCCC 2 | 535719 | 0,595 | 0,030 | 0,154 | 40 | 0,008 | 0,600 | 0,014 | 2,253 | In use
C6000-B:C2 | URIC ACID | PCCC 2-B | 535719 | 0,595 | 0,030 | 0,376 | 33 | 0,019 | 0,606 | 0,012 | 1,987 | In use
C6000-B:C2 | URIC ACID | PCCC 1-B | 525205 | 0,282 | 0,014 | 0,282 | 40 | 0,014 | 0,286 | 0,006 | 2,059 | In use
C6000-B:C2 | URIC ACID URINE | Ran-Liq 3 | 1238UC | 0,850 | 0,080 | 0,363 | 25 | 0,034 | 0,879 | 0,020 | 2,313 | In use
C6000-B:C2 | URIC ACID URINE | Ran-Liq 2 | 1237UC | 0,530 | 0,050 | 0,251 | 25 | 0,024 | 0,543 | 0,017 | 3,154 | In use
Osmometer3320 | Urine osmolality | Ran-Liq 3 | 1238UC | 874,000 | 87,500 | -0,331 | 1 | -0,033 | 845,000 | 0,000 | 0,000 | In use
OSMO1 | Urine osmolality | Ran-Liq 3 | 1238UC | 829,477 | 6,181 | -0,149 | 9 | -0,001 | 828,556 | 5,199 | 0,627 | In use
Osmometer3320 | Urine osmolality | Ran-Liq 2 | 1237UC | 440,000 | 44,000 | 0,068 | 1 | 0,007 | 443,000 | 0,000 | 0,000 | In use
OSMO1 | Urine osmolality | Ran-Liq 2 | 1237UC | 437,063 | 5,905 | -0,434 | 8 | -0,006 | 434,500 | 2,878 | 0,662 | In use
OSMO1 | Urine osmolality | Randox Urine QC 2 | 1237UC | 440,000 | 44,000 | -0,089 | 15 | -0,009 | 436,067 | 3,751 | 0,860 | In use
OSMO1 | Urine osmolality | Randox Urine QC 3 | 1238UC | 874,000 | 87,500 | -0,486 | 13 | -0,049 | 831,462 | 6,565 | 0,790 | In use
C6000-A:E1-2 | VITAMIN B12 II | RIAP L1 | 2179EC | 154,000 | 15,500 | 0,260 | 59 | 0,026 | 158,027 | 9,676 | 6,123 | In use
C6000-A:E1-2 | VITAMIN B12 II | RIAP L2 | 2228EC | 417,000 | 41,500 | 0,581 | 54 | 0,058 | 441,096 | 15,886 | 3,602 | In use

## EFLM targets (hidden)
SEQ | Number | Analyte | Matrix | Within-subject median | Between-subject median | Date Updated on EFLM | Date Scraped | Minimum CV | Minimum Bias | Minimum TEa | Desirable CV | Desirable Bias | Desirable TEa | Optimum CV | Optimum Bias | Optimum TEa
1 | 1 | 17-a-hydroxyprogesterone | Serum/plasma | 28.3 | 38.5 | June 6th 2022 | 2022-06-08 | 21.225 | 17.9183 | 52.9396 | 14.15 | 11.9456 | 35.2931 | 7.075 | 5.9728 | 17.6465
2 | 2 | 25-hydroxy vitamin D3 | Serum/plasma | 6.8 | 26.1 | June 6th 2022 | 2022-06-08 | 5.1 | 10.1142 | 18.5292 | 3.4 | 6.7428 | 12.3528 | 1.7 | 3.3714 | 6.1764
3 | 3 | a-1 antitrypsin | Serum/plasma | 4.1 | 10.5 | June 6th 2022 | 2022-06-08 | 3.075 | 4.227 | 9.3008 | 2.05 | 2.818 | 6.2005 | 1.025 | 1.409 | 3.1003
4 | 4 | Actual bicarbonate | Whole Blood | 4 | 4.8 | February 11th 2021 | 2022-06-08 | 3 | 2.3431 | 7.2931 | 2 | 1.562 | 4.862 | 1 | 0.78102 | 2.431
5 | 5 | Adiponectin | Serum/plasma | 18.8 | 51.2 | October 20th 2019 | 2022-06-08 | 14.1 | 20.4534 | 43.7184 | 9.4 | 13.6356 | 29.1456 | 4.7 | 6.8178 | 14.5728
6 | 6 | Adrenaline | No Matrix Set | 135 | 54 | February 14th 2020 | 2022-06-08 | 101.25 | 54.5248 | 221.5873 | 67.5 | 36.3499 | 147.7249 | 33.75 | 18.1749 | 73.8624
7 | 7 | Alanine transaminase (ALT) | Serum/plasma | 10 | 29.3 | June 6th 2022 | 2022-06-08 | 7.5 | 11.6098 | 23.9848 | 5 | 7.7399 | 15.9899 | 2.5 | 3.8699 | 7.9949
8 | 8 | Albumin | Serum/plasma | 2.5 | 4.9 | June 6th 2022 | 2022-06-08 | 1.875 | 2.0628 | 5.1566 | 1.25 | 1.3752 | 3.4377 | 0.625 | 0.68761 | 1.7189
9 | 9 | Aldosterone | Serum/plasma | 36.6 | 34.7 | October 20th 2019 | 2022-06-08 | 27.45 | 18.913 | 64.2055 | 18.3 | 12.6087 | 42.8037 | 9.15 | 6.3043 | 21.4018
10 | 10 | Alkaline phosphatase (ALP) bone isoenzyme | Serum/plasma | 6.6 | 35.6 | November 9th 2019 | 2022-06-08 | 4.95 | 13.5775 | 21.745 | 3.3 | 9.0517 | 14.4967 | 1.65 | 4.5258 | 7.2483
11 | 11 | Alkaline phosphatase (ALP), liver type | Serum/plasma | 5.3 | 24 | June 6th 2022 | 2022-06-08 | 3.975 | 9.2168 | 15.7756 | 2.65 | 6.1446 | 10.5171 | 1.325 | 3.0723 | 5.2585
12 | 12 | Alphafetoprotein | Serum/plasma | 4.5 | 55.1 | April 19th 2022 | 2022-06-08 | 3.375 | 20.7313 | 26.3 | 2.25 | 13.8209 | 17.5334 | 1.125 | 6.9104 | 8.7667
13 | 13 | Amylase | Serum/plasma | 6.5 | 30.2 | June 6th 2022 | 2022-06-08 | 4.875 | 11.5843 | 19.6281 | 3.25 | 7.7229 | 13.0854 | 1.625 | 3.8614 | 6.5427
14 | 14 | Antimullerian hormone | Serum/plasma | 19.2 |  | December 14th 2020 | 2022-06-08 | 14.4 | 7.2 | 30.96 | 9.6 | 4.8 | 20.64 | 4.8 | 2.4 | 10.32
15 | 15 | Apolipoprotein A1 | Serum/plasma | 5.3 | 11.1 | June 6th 2022 | 2022-06-08 | 3.975 | 4.6127 | 11.1714 | 2.65 | 3.0751 | 7.4476 | 1.325 | 1.5376 | 3.7238
16 | 16 | Apolipoprotein B | Serum/plasma | 7.3 | 20.1 | June 6th 2022 | 2022-06-08 | 5.475 | 8.0192 | 17.053 | 3.65 | 5.3461 | 11.3686 | 1.825 | 2.6731 | 5.6843
17 | 17 | Aspartate transaminase (AST) | Serum/plasma | 9.5 | 20.7 | June 6th 2022 | 2022-06-08 | 7.125 | 8.5409 | 20.2972 | 4.75 | 5.694 | 13.5315 | 2.375 | 2.847 | 6.7657
18 | 18 | b-2 microglobulin | Serum/plasma | 4.1 | 11.5 | June 6th 2022 | 2022-06-08 | 3.075 | 4.5784 | 9.6521 | 2.05 | 3.0523 | 6.4348 | 1.025 | 1.5261 | 3.2174
19 | 19 | Basophils | Whole Blood | 12.4 | 26.3 | June 6th 2022 | 2022-06-08 | 9.3 | 10.9037 | 26.2487 | 6.2 | 7.2692 | 17.4992 | 3.1 | 3.6346 | 8.7496
20 | 20 | Bilirubin | Serum/plasma | 20 | 26.6 | June 6th 2022 | 2022-06-08 | 15 | 12.48 | 37.23 | 10 | 8.32 | 24.82 | 5 | 4.16 | 12.41
21 | 21 | Calcitonin | Serum/plasma | 13 | 65.8 | January 14th 2022 | 2022-06-08 | 9.75 | 25.152 | 41.2395 | 6.5 | 16.768 | 27.493 | 3.25 | 8.384 | 13.7465
22 | 22 | Calcium (Ca) | Serum/plasma | 1.8 | 2.7 | June 6th 2022 | 2022-06-08 | 1.35 | 1.2169 | 3.4444 | 0.9 | 0.81125 | 2.2962 | 0.45 | 0.40562 | 1.1481
23 | 23 | Cancer antigen 125 (CA-125) | Serum/plasma | 8.7 | 25.4 | June 6th 2022 | 2022-06-08 | 6.525 | 10.0682 | 20.8345 | 4.35 | 6.7122 | 13.8897 | 2.175 | 3.3561 | 6.9448
24 | 24 | Cancer antigen 19-9 (CA 19-9) | Serum/plasma | 4.3 | 57.4 | June 6th 2022 | 2022-06-08 | 3.225 | 21.5853 | 26.9066 | 2.15 | 14.3902 | 17.9377 | 1.075 | 7.1951 | 8.9689
25 | 25 | Cancer antigen 72-4 (CA 72-4) | Serum/plasma | 50.3 | 103.4 | November 20th 2019 | 2022-06-08 | 37.725 | 43.1195 | 105.3658 | 25.15 | 28.7464 | 70.2439 | 12.575 | 14.3732 | 35.1219
26 | 26 | Carbon dioxide (total) | Whole Blood | 4 | 4.8 | February 11th 2021 | 2022-06-08 | 3 | 2.3431 | 7.2931 | 2 | 1.562 | 4.862 | 1 | 0.78102 | 2.431
27 | 27 | Carcinoembryonic antigen (CEA) | Serum/plasma | 6.8 | 59.3 | June 6th 2022 | 2022-06-08 | 5.1 | 22.3832 | 30.7982 | 3.4 | 14.9222 | 20.5322 | 1.7 | 7.4611 | 10.2661
28 | 28 | Ceruloplasmin | Serum/plasma | 4.9 | 15.2 | May 26th 2020 | 2022-06-08 | 3.675 | 5.9889 | 12.0526 | 2.45 | 3.9926 | 8.0351 | 1.225 | 1.9963 | 4.0175
29 | 29 | Chloride | Serum/plasma | 1 | 1.3 | June 6th 2022 | 2022-06-08 | 0.75 | 0.61505 | 1.8525 | 0.5 | 0.41003 | 1.235 | 0.25 | 0.20502 | 0.61752
30 | 30 | Cholesterol | Serum/plasma | 5.3 | 16.2 | June 6th 2022 | 2022-06-08 | 3.975 | 6.3919 | 12.9506 | 2.65 | 4.2612 | 8.6337 | 1.325 | 2.1306 | 4.3169
31 | 31 | Chromogranin A | Serum/plasma | 14.8 | 30.6 | June 6th 2022 | 2022-06-08 | 11.1 | 12.7467 | 31.0617 | 7.4 | 8.4978 | 20.7078 | 3.7 | 4.2489 | 10.3539
32 | 32 | Complement 3 (C3) | Serum/plasma | 4.6 | 15.2 | November 25th 2019 | 2022-06-08 | 3.45 | 5.9553 | 11.6478 | 2.3 | 3.9702 | 7.7652 | 1.15 | 1.9851 | 3.8826
33 | 33 | Complement 4 (C4) | Serum/plasma | 6.9 | 24.5 | April 29th 2020 | 2022-06-08 | 5.175 | 9.5449 | 18.0837 | 3.45 | 6.3633 | 12.0558 | 1.725 | 3.1816 | 6.0279
34 | 34 | Copper | Serum/plasma | 7.4 | 13.2 | June 6th 2022 | 2022-06-08 | 5.55 | 5.6748 | 14.8323 | 3.7 | 3.7832 | 9.8882 | 1.85 | 1.8916 | 4.9441
35 | 35 | Cortisol | Serum/plasma | 16.3 | 48.7 | June 6th 2022 | 2022-06-08 | 12.225 | 19.2583 | 39.4295 | 8.15 | 12.8389 | 26.2864 | 4.075 | 6.4194 | 13.1432
36 | 36 | C-reactive protein (CRP) | Serum/plasma | 34 | 83.6 | June 6th 2022 | 2022-06-08 | 25.5 | 33.8435 | 75.9185 | 17 | 22.5624 | 50.6124 | 8.5 | 11.2812 | 25.3062
37 | 37 | C-reactive protein (CRP) - high-sensitive | Serum/plasma | 58.9 | 77.4 | January 12th 2020 | 2022-06-08 | 44.175 | 36.4734 | 109.3621 | 29.45 | 24.3156 | 72.9081 | 14.725 | 12.1578 | 36.454
38 | 38 | Creatine Kinase | Serum/plasma | 15 | 37.9 | June 6th 2022 | 2022-06-08 | 11.25 | 15.2851 | 33.8476 | 7.5 | 10.1901 | 22.5651 | 3.75 | 5.095 | 11.2825
39 | 39 | Creatinine | Serum/plasma | 4.4 | 14.1 | June 6th 2022 | 2022-06-08 | 3.3 | 5.539 | 10.984 | 2.2 | 3.6926 | 7.3226 | 1.1 | 1.8463 | 3.6613
40 | 40 | Cystatin C | Serum/plasma | 3.9 | 12.1 | June 6th 2022 | 2022-06-08 | 2.925 | 4.7674 | 9.5936 | 1.95 | 3.1782 | 6.3957 | 0.975 | 1.5891 | 3.1979
41 | 41 | Dehydroepiandrosterone sulphate (DHEAS) | Serum/plasma | 5.9 | 21 | June 6th 2022 | 2022-06-08 | 4.425 | 8.1799 | 15.4812 | 2.95 | 5.4533 | 10.3208 | 1.475 | 2.7266 | 5.1604
42 | 42 | Dihydroxytestosterone (DHT) | Serum/plasma | 11.5 | 37.1 | June 6th 2022 | 2022-06-08 | 8.625 | 14.5656 | 28.7968 | 5.75 | 9.7104 | 19.1979 | 2.875 | 4.8552 | 9.5989
43 | 43 | Eosinophils | Whole Blood | 14.9 | 65.3 | June 6th 2022 | 2022-06-08 | 11.175 | 25.1169 | 43.5556 | 7.45 | 16.7446 | 29.0371 | 3.725 | 8.3723 | 14.5185
44 | 44 | Erythrocytes | Whole Blood | 2.6 | 6.4 | June 6th 2022 | 2022-06-08 | 1.95 | 2.5905 | 5.808 | 1.3 | 1.727 | 3.872 | 0.65 | 0.8635 | 1.936
45 | 45 | Estradiol | Serum/plasma | 15 | 13 | April 20th 2020 | 2022-06-08 | 11.25 | 7.4435 | 26.006 | 7.5 | 4.9624 | 17.3374 | 3.75 | 2.4812 | 8.6687
46 | 46 | Ferritin | Serum/plasma | 12.8 |  | July 2nd 2020 | 2022-06-08 | 9.6 | 4.8 | 20.64 | 6.4 | 3.2 | 13.76 | 3.2 | 1.6 | 6.88
47 | 47 | Fibroblast growth factor-23 | Serum/plasma | 13.9 | 21.4 | June 6th 2022 | 2022-06-08 | 10.425 | 9.5693 | 26.7705 | 6.95 | 6.3795 | 17.847 | 3.475 | 3.1898 | 8.9235
48 | 48 | Follicle stimulating hormone (FSH) | Serum/plasma | 12.3 | 42.1 | June 6th 2022 | 2022-06-08 | 9.225 | 16.4475 | 31.6688 | 6.15 | 10.965 | 21.1125 | 3.075 | 5.4825 | 10.5563
49 | 49 | Free kappa light chain | Serum/plasma | 4.8 | 15.3 | May 26th 2020 | 2022-06-08 | 3.6 | 6.0132 | 11.9532 | 2.4 | 4.0088 | 7.9688 | 1.2 | 2.0044 | 3.9844
50 | 50 | Free lambda light chain | Serum/plasma | 4.8 | 17.3 | November 20th 2019 | 2022-06-08 | 3.6 | 6.7326 | 12.6726 | 2.4 | 4.4884 | 8.4484 | 1.2 | 2.2442 | 4.2242
51 | 51 | Fructosamine | Serum/plasma | 2.3 | 6.3 | November 20th 2019 | 2022-06-08 | 1.725 | 2.515 | 5.3613 | 1.15 | 1.6767 | 3.5742 | 0.575 | 0.83834 | 1.7871
52 | 52 | g-glutamyl transferase activity | Serum/plasma | 9.1 | 44.4 | June 6th 2022 | 2022-06-08 | 6.825 | 16.9961 | 28.2574 | 4.55 | 11.3307 | 18.8382 | 2.275 | 5.6654 | 9.4191
53 | 53 | Glucose | Serum/plasma | 4.9 | 8.1 | June 6th 2022 | 2022-06-08 | 3.675 | 3.55 | 9.6138 | 2.45 | 2.3667 | 6.4092 | 1.225 | 1.1833 | 3.2046
54 | 54 | Glycated albumin | Serum/plasma | 1.4 | 5.7 | June 6th 2022 | 2022-06-08 | 1.05 | 2.201 | 3.9335 | 0.7 | 1.4674 | 2.6224 | 0.35 | 0.73368 | 1.3112
55 | 55 | Haematocrit (Hct) | Whole Blood | 2.8 | 5.4 | June 6th 2022 | 2022-06-08 | 2.1 | 2.281 | 5.746 | 1.4 | 1.5207 | 3.8307 | 0.7 | 0.76035 | 1.9153
56 | 56 | Haemoglobin A1c (IFCC) | Whole Blood | 1.6 | 7.1 | June 6th 2022 | 2022-06-08 | 1.2 | 2.7293 | 4.7093 | 0.8 | 1.8195 | 3.1395 | 0.4 | 0.90976 | 1.5698
57 | 57 | Haemoglobin A1c (NGSP) | Whole Blood | 1.1 | 4.8 | June 6th 2022 | 2022-06-08 | 0.825 | 1.8467 | 3.2079 | 0.55 | 1.2311 | 2.1386 | 0.275 | 0.61555 | 1.0693
58 | 58 | Haemoglobin (Hb) | Whole Blood | 2.7 | 5.9 | June 6th 2022 | 2022-06-08 | 2.025 | 2.4332 | 5.7744 | 1.35 | 1.6221 | 3.8496 | 0.675 | 0.81106 | 1.9248
59 | 59 | Haptoglobin | Serum/plasma | 8.6 | 39 | June 6th 2022 | 2022-06-08 | 6.45 | 14.9764 | 25.6189 | 4.3 | 9.9842 | 17.0792 | 2.15 | 4.9921 | 8.5396
60 | 60 | HDL cholesterol | Serum/plasma | 5.7 | 24.4 | June 6th 2022 | 2022-06-08 | 4.275 | 9.3964 | 16.4501 | 2.85 | 6.2642 | 10.9667 | 1.425 | 3.1321 | 5.4834
61 | 61 | Human epididymis protein 4 (HE4) | Serum/plasma | 7 | 17.9 | April 19th 2022 | 2022-06-08 | 5.25 | 7.2075 | 15.87 | 3.5 | 4.805 | 10.58 | 1.75 | 2.4025 | 5.29
62 | 62 | Hydrogen ion (H+) | Whole Blood | 3.5 | 2 | February 11th 2021 | 2022-06-08 | 2.625 | 1.5117 | 5.8429 | 1.75 | 1.0078 | 3.8953 | 0.875 | 0.50389 | 1.9476
63 | 63 | Immunoglobulin A (IgA) | Serum/plasma | 5.6 | 19.5 | June 6th 2022 | 2022-06-08 | 4.2 | 7.6081 | 14.5381 | 2.8 | 5.072 | 9.692 | 1.4 | 2.536 | 4.846
64 | 64 | Immunoglobulin G (IgG) | Serum/plasma | 3.4 | 17.1 | June 6th 2022 | 2022-06-08 | 2.55 | 6.538 | 10.7455 | 1.7 | 4.3587 | 7.1637 | 0.85 | 2.1793 | 3.5818
65 | 65 | Immunoglobulin M (IgM) | Serum/plasma | 5.9 | 48.5 | January 10th 2020 | 2022-06-08 | 4.425 | 18.3216 | 25.6228 | 2.95 | 12.2144 | 17.0819 | 1.475 | 6.1072 | 8.5409
66 | 66 | Insulin | Serum/plasma | 25.4 | 33.4 | June 6th 2022 | 2022-06-08 | 19.05 | 15.7354 | 47.1679 | 12.7 | 10.4902 | 31.4452 | 6.35 | 5.2451 | 15.7226
67 | 67 | Insulin-like growth factor-1 (IGF-1) | Serum/plasma | 9.4 | 27 | October 29th 2019 | 2022-06-08 | 7.05 | 10.7211 | 22.3536 | 4.7 | 7.1474 | 14.9024 | 2.35 | 3.5737 | 7.4512
68 | 68 | Interleukin-6 (IL6) | Serum/plasma | 27.8 | 33.9 | June 6th 2022 | 2022-06-08 | 20.85 | 16.4404 | 50.8429 | 13.9 | 10.9603 | 33.8953 | 6.95 | 5.4801 | 16.9476
69 | 69 | Iron | Serum/plasma | 20.6 | 32.3 | June 6th 2022 | 2022-06-08 | 15.45 | 14.3662 | 39.8587 | 10.3 | 9.5775 | 26.5725 | 5.15 | 4.7887 | 13.2862
70 | 70 | Lactate dehydrogenase (LDH) activity | Serum/plasma | 5.2 | 12.6 | November 18th 2020 | 2022-06-08 | 3.9 | 5.1116 | 11.5466 | 2.6 | 3.4077 | 7.6977 | 1.3 | 1.7039 | 3.8489
71 | 71 | LDL Cholesterol | Serum/plasma | 8.3 | 26.1 | June 6th 2022 | 2022-06-08 | 6.225 | 10.2705 | 20.5417 | 4.15 | 6.847 | 13.6945 | 2.075 | 3.4235 | 6.8472
72 | 72 | Leukocytes | Whole Blood | 10.8 | 16.4 | June 6th 2022 | 2022-06-08 | 8.1 | 7.3638 | 20.7288 | 5.4 | 4.9092 | 13.8192 | 2.7 | 2.4546 | 6.9096
73 | 73 | Lipase | Serum/plasma | 9.2 | 24.8 | June 6th 2022 | 2022-06-08 | 6.9 | 9.9193 | 21.3043 | 4.6 | 6.6129 | 14.2029 | 2.3 | 3.3064 | 7.1014
74 | 74 | Lipoprotein (a) | Serum/plasma | 8.8 |  | January 13th 2022 | 2022-06-08 | 6.6 | 3.3 | 14.19 | 4.4 | 2.2 | 9.46 | 2.2 | 1.1 | 4.73
75 | 75 | Luteinising hormone (LH) | Serum/plasma | 22.7 | 30.8 | June 6th 2022 | 2022-06-08 | 17.025 | 14.348 | 42.4393 | 11.35 | 9.5653 | 28.2928 | 5.675 | 4.7827 | 14.1464
76 | 76 | Lymphocytes | Whole Blood | 10.8 | 22.6 | June 6th 2022 | 2022-06-08 | 8.1 | 9.393 | 22.758 | 5.4 | 6.262 | 15.172 | 2.7 | 3.131 | 7.586
77 | 77 | Magnesium | Serum/plasma | 2.8 | 5.7 | June 6th 2022 | 2022-06-08 | 2.1 | 2.3815 | 5.8465 | 1.4 | 1.5876 | 3.8976 | 0.7 | 0.79382 | 1.9488
79 | 78 | Mean corpuscular haemoglobin concentration (MCHC) | Whole Blood | 0.9 | 1.4 | June 6th 2022 | 2022-06-08 | 0.675 | 0.62412 | 1.7379 | 0.45 | 0.41608 | 1.1586 | 0.225 | 0.20804 | 0.57929
78 | 79 | Mean corpuscular haemoglobin (MCH) | Whole Blood | 0.8 | 4.3 | June 6th 2022 | 2022-06-08 | 0.6 | 1.6402 | 2.6302 | 0.4 | 1.0934 | 1.7534 | 0.2 | 0.54672 | 0.87672
80 | 80 | Mean corpuscular volume (MCV) | Whole Blood | 0.8 | 3.7 | June 6th 2022 | 2022-06-08 | 0.6 | 1.4196 | 2.4096 | 0.4 | 0.94637 | 1.6064 | 0.2 | 0.47319 | 0.80319
81 | 81 | Mean platelet volume (MPV) | Whole Blood | 2.2 | 7 | June 6th 2022 | 2022-06-08 | 1.65 | 2.7516 | 5.4741 | 1.1 | 1.8344 | 3.6494 | 0.55 | 0.9172 | 1.8247
82 | 82 | Monocytes | Whole Blood | 13.3 | 22.2 | June 6th 2022 | 2022-06-08 | 9.975 | 9.7047 | 26.1634 | 6.65 | 6.4698 | 17.4423 | 3.325 | 3.2349 | 8.7211
83 | 83 | Neuron specific enolase | Serum/plasma | 10.8 | 16.6 | June 6th 2022 | 2022-06-08 | 8.1 | 7.4265 | 20.7915 | 5.4 | 4.951 | 13.861 | 2.7 | 2.4755 | 6.9305
84 | 84 | Neutrophils | Whole Blood | 14 | 23.5 | June 6th 2022 | 2022-06-08 | 10.5 | 10.2578 | 27.5828 | 7 | 6.8385 | 18.3885 | 3.5 | 3.4193 | 9.1943
85 | 85 | Orosmucoid | Serum/plasma | 7.3 | 24 | June 6th 2022 | 2022-06-08 | 5.475 | 9.4071 | 18.4409 | 3.65 | 6.2714 | 12.2939 | 1.825 | 3.1357 | 6.147
86 | 86 | Osteocalcin | Serum/plasma | 8.9 | 32.3 | June 6th 2022 | 2022-06-08 | 6.675 | 12.5639 | 23.5776 | 4.45 | 8.3759 | 15.7184 | 2.225 | 4.188 | 7.8592
87 | 87 | Pancreatic Amylase | Serum/plasma | 6.6 | 25.7 | June 6th 2022 | 2022-06-08 | 4.95 | 9.9502 | 18.1177 | 3.3 | 6.6335 | 12.0785 | 1.65 | 3.3167 | 6.0392
88 | 88 | Parathyroid hormone | Serum/plasma | 15.7 | 23.5 | June 6th 2022 | 2022-06-08 | 11.775 | 10.5982 | 30.027 | 7.85 | 7.0655 | 20.018 | 3.925 | 3.5327 | 10.009
89 | 89 | Partial pressure carbon dioxide | Whole Blood | 4.8 | 5.3 | February 11th 2021 | 2022-06-08 | 3.6 | 2.6814 | 8.6214 | 2.4 | 1.7876 | 5.7476 | 1.2 | 0.89382 | 2.8738
90 | 90 | Phosphate | Serum/plasma | 7.7 | 10.7 | June 6th 2022 | 2022-06-08 | 5.775 | 4.9435 | 14.4722 | 3.85 | 3.2956 | 9.6481 | 1.925 | 1.6478 | 4.8241
93 | 91 | Plateletcrit | Whole Blood | 6.4 | 13.5 | June 6th 2022 | 2022-06-08 | 4.8 | 5.6026 | 13.5226 | 3.2 | 3.7351 | 9.0151 | 1.6 | 1.8675 | 4.5075
91 | 92 | Platelet distribution wide (PDW) | Whole Blood | 3.8 | 12.3 | June 6th 2022 | 2022-06-08 | 2.85 | 4.8276 | 9.5301 | 1.9 | 3.2184 | 6.3534 | 0.95 | 1.6092 | 3.1767
92 | 93 | Platelet larger cell ratio (P-LCR) | Whole Blood | 6.7 | 20.9 | July 3rd 2020 | 2022-06-08 | 5.025 | 8.2304 | 16.5216 | 3.35 | 5.4869 | 11.0144 | 1.675 | 2.7435 | 5.5072
94 | 94 | Potassium | Serum/plasma | 4 | 4.1 | June 6th 2022 | 2022-06-08 | 3 | 2.148 | 7.098 | 2 | 1.432 | 4.732 | 1 | 0.716 | 2.366
95 | 95 | Prolactin | Serum/plasma | 29.5 | 43 | June 6th 2022 | 2022-06-08 | 22.125 | 19.5549 | 56.0612 | 14.75 | 13.0366 | 37.3741 | 7.375 | 6.5183 | 18.6871
96 | 96 | Prostate specific antigen (PSA) | Serum/plasma | 6.8 | 42 | October 20th 2019 | 2022-06-08 | 5.1 | 15.9551 | 24.3701 | 3.4 | 10.6367 | 16.2467 | 1.7 | 5.3184 | 8.1234
97 | 97 | Prostate specific antigen (PSA) - Conjugated | Serum/plasma | 8.8 | 57.7 | November 9th 2019 | 2022-06-08 | 6.6 | 21.8877 | 32.7777 | 4.4 | 14.5918 | 21.8518 | 2.2 | 7.2959 | 10.9259
98 | 98 | Prostate specific antigen (PSA) - free | Serum/plasma | 7.1 | 46.2 | October 20th 2019 | 2022-06-08 | 5.325 | 17.5284 | 26.3146 | 3.55 | 11.6856 | 17.5431 | 1.775 | 5.8428 | 8.7715
99 | 99 | Protein (total) | Serum/plasma | 2.6 | 4.6 | June 6th 2022 | 2022-06-08 | 1.95 | 1.9815 | 5.199 | 1.3 | 1.321 | 3.466 | 0.65 | 0.66049 | 1.733
100 | 100 | Red cell distribution wide standard deviation (RDW-SD) | Whole Blood | 1.6 | 4.2 | June 6th 2022 | 2022-06-08 | 1.2 | 1.6854 | 3.6654 | 0.8 | 1.1236 | 2.4436 | 0.4 | 0.56181 | 1.2218
101 | 101 | Renin | Serum/plasma | 30.1 | 41.6 | November 18th 2019 | 2022-06-08 | 22.575 | 19.2553 | 56.5041 | 15.05 | 12.8369 | 37.6694 | 7.525 | 6.4184 | 18.8347
102 | 102 | Reticulocyte haemoglobin equivalent (RET-He) | Whole Blood | 1.7 | 3.3 | June 6th 2022 | 2022-06-08 | 1.275 | 1.3921 | 3.4958 | 0.85 | 0.92804 | 2.3305 | 0.425 | 0.46402 | 1.1653
103 | 103 | Reticulocytes | Whole Blood | 9.7 | 27.1 | June 6th 2022 | 2022-06-08 | 7.275 | 10.7939 | 22.7976 | 4.85 | 7.1959 | 15.1984 | 2.425 | 3.598 | 7.5992
104 | 104 | S100 calcium-binding protein B (S100B) | Serum/plasma | 10.2 | 32.7 | November 20th 2019 | 2022-06-08 | 7.65 | 12.8452 | 25.4677 | 5.1 | 8.5635 | 16.9785 | 2.55 | 4.2817 | 8.4892
105 | 105 | Selenium | Serum/plasma | 7.7 | 11.3 | June 6th 2022 | 2022-06-08 | 5.775 | 5.1278 | 14.6565 | 3.85 | 3.4185 | 9.771 | 1.925 | 1.7093 | 4.8855
106 | 106 | Sexual-hormone-binding-globulin (SHBG) | Serum/plasma | 9.6 | 35.5 | June 6th 2022 | 2022-06-08 | 7.2 | 13.7907 | 25.6707 | 4.8 | 9.1938 | 17.1138 | 2.4 | 4.5969 | 8.5569
107 | 107 | Sodium | Serum/plasma | 0.5 | 0.9 | June 6th 2022 | 2022-06-08 | 0.375 | 0.38609 | 1.0048 | 0.25 | 0.25739 | 0.66989 | 0.125 | 0.1287 | 0.33495
108 | 108 | Soluble transferrin receptor (sTfr) | Serum/plasma | 6.9 | 19.1 | May 26th 2020 | 2022-06-08 | 5.175 | 7.6155 | 16.1543 | 3.45 | 5.077 | 10.7695 | 1.725 | 2.5385 | 5.3848
109 | 109 | Standard bicarbonate | Serum/plasma | 4.2 | 4.4 | February 11th 2021 | 2022-06-08 | 3.15 | 2.281 | 7.4785 | 2.1 | 1.5207 | 4.9857 | 1.05 | 0.76035 | 2.4928
110 | 110 | Suppression of tumorigenicity 2 (ST2) | Serum/plasma | 10.5 | 30.4 | November 20th 2019 | 2022-06-08 | 7.875 | 12.0608 | 25.0546 | 5.25 | 8.0406 | 16.7031 | 2.625 | 4.0203 | 8.3515
111 | 111 | Tartrate-resistant acid phosphatase (TRAP) | Serum/plasma | 10.8 | 13.3 | November 25th 2019 | 2022-06-08 | 8.1 | 6.4248 | 19.7898 | 5.4 | 4.2832 | 13.1932 | 2.7 | 2.1416 | 6.5966
112 | 112 | Testosterone | Serum/plasma | 12.5 | 21.4 | June 6th 2022 | 2022-06-08 | 9.375 | 9.2937 | 24.7625 | 6.25 | 6.1958 | 16.5083 | 3.125 | 3.0979 | 8.2542
113 | 113 | Testosterone - free | Serum/plasma | 21.9 | 29 | June 6th 2022 | 2022-06-08 | 16.425 | 13.6276 | 40.7288 | 10.95 | 9.085 | 27.1525 | 5.475 | 4.5425 | 13.5763
114 | 114 | Thrombocytes | Whole Blood | 7.5 | 18.6 | June 6th 2022 | 2022-06-08 | 5.625 | 7.5207 | 16.8019 | 3.75 | 5.0138 | 11.2013 | 1.875 | 2.5069 | 5.6006
115 | 115 | Thyroglobulin | Serum/plasma | 10.5 | 77.2 | June 6th 2022 | 2022-06-08 | 7.875 | 29.2165 | 42.2103 | 5.25 | 19.4777 | 28.1402 | 2.625 | 9.7388 | 14.0701
116 | 116 | Thyroid stimulating hormone (TSH) | Serum/plasma | 17.6 | 35.9 | June 6th 2022 | 2022-06-08 | 13.2 | 14.9933 | 36.7733 | 8.8 | 9.9955 | 24.5155 | 4.4 | 4.9978 | 12.2578
117 | 117 | Thyroxine - free (FT4) | Serum/plasma | 4.8 | 7.7 | June 6th 2022 | 2022-06-08 | 3.6 | 3.4026 | 9.3426 | 2.4 | 2.2684 | 6.2284 | 1.2 | 1.1342 | 3.1142
118 | 118 | Thyroxine - total (T4) | Serum/plasma | 6.4 | 12 | June 6th 2022 | 2022-06-08 | 4.8 | 5.1 | 13.02 | 3.2 | 3.4 | 8.68 | 1.6 | 1.7 | 4.34
119 | 119 | Total tri-iodothyronine (T3) | Serum/plasma | 9.4 | 12.2 | June 6th 2022 | 2022-06-08 | 7.05 | 5.7755 | 17.408 | 4.7 | 3.8503 | 11.6053 | 2.35 | 1.9252 | 5.8027
120 | 120 | Transferrin | Serum/plasma | 3.8 | 13.9 | June 6th 2022 | 2022-06-08 | 2.85 | 5.4038 | 10.1063 | 1.9 | 3.6025 | 6.7375 | 0.95 | 1.8013 | 3.3688
121 | 121 | Transthyretin | Serum/plasma | 10.9 | 19.1 | January 10th 2020 | 2022-06-08 | 8.175 | 8.2468 | 21.7355 | 5.45 | 5.4978 | 14.4903 | 2.725 | 2.7489 | 7.2452
122 | 122 | Triglycerides | Serum/plasma | 20 | 37 | June 6th 2022 | 2022-06-08 | 15 | 15.7723 | 40.5223 | 10 | 10.5149 | 27.0149 | 5 | 5.2574 | 13.5074
123 | 123 | Triiodothyronine - free (FT3) | Serum/plasma | 4.9 | 8.3 | June 6th 2022 | 2022-06-08 | 3.675 | 3.6144 | 9.6782 | 2.45 | 2.4096 | 6.4521 | 1.225 | 1.2048 | 3.2261
124 | 124 | Troponin I, cardiac - high-sensitive (biweekly-monthly sampling) | Serum/plasma | 12 | 35.8 | June 6th 2022 | 2022-06-08 | 9 | 14.1591 | 29.0091 | 6 | 9.4394 | 19.3394 | 3 | 4.7197 | 9.6697
125 | 125 | Troponin T, cardiac - high-sensitive (biweekly-monthly sampling) | Serum/plasma | 11.3 | 30.8 | June 6th 2022 | 2022-06-08 | 8.475 | 12.3028 | 26.2866 | 5.65 | 8.2019 | 17.5244 | 2.825 | 4.1009 | 8.7622
126 | 126 | Tumor necrosis factor (TNF) | Serum/plasma | 16.1 | 26.7 | June 6th 2022 | 2022-06-08 | 12.075 | 11.6919 | 31.6157 | 8.05 | 7.7946 | 21.0771 | 4.025 | 3.8973 | 10.5386
127 | 127 | Tyrosine | Serum/plasma | 7.5 | 11.9 | November 20th 2019 | 2022-06-08 | 5.625 | 5.2749 | 14.5561 | 3.75 | 3.5166 | 9.7041 | 1.875 | 1.7583 | 4.852
128 | 128 | Urate | Serum/plasma | 8.2 | 22.4 | June 6th 2022 | 2022-06-08 | 6.15 | 8.9451 | 19.0926 | 4.1 | 5.9634 | 12.7284 | 2.05 | 2.9817 | 6.3642
129 | 129 | Urea | Serum/plasma | 13.9 | 20.9 | June 6th 2022 | 2022-06-08 | 10.425 | 9.4126 | 26.6138 | 6.95 | 6.275 | 17.7425 | 3.475 | 3.1375 | 8.8713
130 | 130 | Vitamin A | Serum/plasma | 6.1 | 21 | June 6th 2022 | 2022-06-08 | 4.575 | 8.2005 | 15.7493 | 3.05 | 5.467 | 10.4995 | 1.525 | 2.7335 | 5.2498
131 | 131 | Vitamin E | Serum/plasma | 7 | 17.3 | June 6th 2022 | 2022-06-08 | 5.25 | 6.9984 | 15.6609 | 3.5 | 4.6656 | 10.4406 | 1.75 | 2.3328 | 5.2203
132 | 132 | Zinc | Serum/plasma | 8.5 | 8.8 | June 6th 2022 | 2022-06-08 | 6.375 | 4.588 | 15.1068 | 4.25 | 3.0587 | 10.0712 | 2.125 | 1.5293 | 5.0356

## Diasorin
Instrument | Test | QC | QC lot No. | Target mean | Target SD | SDI | Number of values | Bias | Calculated mean | Calculated SD | CV% | Usage
LIAISON XL | ALDO | ALDO L1 | 229352 | 178,390 | 17,820 | 1,227 | 8 | 0,123 | 200,250 | 16,816 | 8,398 | In use
LIAISON XL | ALDO | ALDO L2 | 230352 | 918,235 | 91,878 | 0,445 | 7 | 0,045 | 959,143 | 39,541 | 4,123 | In use
LIAISON XL | Human growth hormone | hGH L1 | 7980320 | 3,700 | 0,247 | 0,939 | 6 | 0,063 | 3,932 | 0,056 | 1,415 | In use
LIAISON XL | Human growth hormone | hGH L2 | 9260320 | 17,350 | 1,150 | 1,275 | 6 | 0,085 | 18,817 | 0,387 | 2,056 | In use
LIAISON XL | Insulin-like growth factor I | IGF-1 L1 | 1412471 | 88,850 | 9,383 | 0,724 | 4 | 0,076 | 95,643 | 7,189 | 7,516 | In use
LIAISON XL | Insulin-like growth factor I | IGF-1 L2 | 1422471 | 341,500 | 29,167 | -0,267 | 4 | -0,023 | 333,700 | 25,222 | 7,558 | In use
LIAISON XL | Renin | RENIN L1 | 7520333 | 27,785 | 2,315 | -0,850 | 5 | -0,071 | 25,818 | 0,945 | 3,661 | In use
LIAISON XL | Renin | RENIN L2 | 7530333 | 109,460 | 9,113 | -0,073 | 5 | -0,006 | 108,792 | 7,030 | 6,462 | In use